9 Aug 2007

NCSX

ROGOWSKI COIL

HIGH TEMPERATURE

FINAL DESIGN REVIEW

# Requirements

1.	Provide a robust design and locate in secure location on the outer surface of the vacuum vessel ( Field Period Spacer is an alternate).

It is inaccessible after the field periods are fully constructed.

2.     	Install 3 Coils located approximately 120 ° apart.

3.	Locate as close to the plasma as possible

4.	Survive bake temperatures of 350ºC

- NA of the NCSX greater than 2 sq inches per inch of length

6.	Minimize the gap between the coil centers at the mandrel ends. 3/8 inch is an upper limit.

7.	Fix the open ends and line up the mandrel ends.    Permit the differential thermal expansion to deflect the coil off the surface of the vacuum vessel.

8.	Twist the exiting lead wires immediately outside the footprint of the mandrel and exit the cryostat at the Port 12 horse collar through a Flux Loop junction box

- An upper response in the 1 kHz range. Determine response by measurement of a representative prototype.

- Provide electrostatic shield

- Perform thermal cycle test of a prototype.

- Limit the radial build to a maximum of one inch from the vacuum vessel surface

## Design

- Located on vacuum vessel surface adjacent to weld flange.

Length is approximately 10 ft. Determine each length by measurement

- Mandrel 3/8 inch OD with 0.010 inch wall annealed 316 SS furnished as straight tube to facilitate winding.  μ < 1.02 ( measured )

Three times vendor minimum bend radius of co ax  cable

- Wire is 0.032 inch diameter coaxial mineral insulated cable, annealed 316 SS conductor and sheath, and compressed MgO insulator.                    μ < 1.02 ( Measured )

Allowable temperature 900ºC in air

- 30 turns per inch ( theoretical max is 31 )

NA = 3.9 inch² per inch  ( A=0.130 inch² )

5.	One layer of Nextel tape 0.018 inch thick,

zero lap wrap

Rated continuous temperature 1100°C

6.        Protective outer housing annealed 316 SS, 5/8 OD by ½ inch ID. μ < 1.02 ( Measured )

- Return wire run in ID of mandrel with Nextel  sleeve to damp vibration

8.	Zirconia spacer at mandrel to mandrel end with flat  machined to accommodate run of wire inside of mandrel

9      Wind co axial cable on mandrel at PPPL

Work station 80 % complete TFTR TC basement

10	 Hold in place with 0.080 inch thick by ½ inch wide inconel 625 formed clamp and skip weld to vessel.  4 clamps, custom bend, trim and weld.

11.	Route 20 ft exiting twisted leads to convenient Flux Loops Junction Box, strain relieve and terminate to upper ( smaller ) circuit board terminal blocks

12.	 0.180 inch differential thermal expansion between the coil outer protective sheath and the vessel. Permit Rogowski Coil to deflect off vessel surface symmetrically  from fixed location   ( Zirconia spacer )

Two additional Inconel clamps sized to permit the differential thermal expansion ( 0.010 inch clearance )

13    Twist coaxial cable leads with 5/8 to ¾ inch pitch

14     One Rogowski to be operational for first plasma.

Conventional cable run from junction box  outside cryostat to electronics  rack ( by others )

- Where  Rogowski Coil is routed over Flux Loops or Voltage Loops provide 316 SS wear plate of 0.020 inch shim stock and spot weld to vessel

16.	Construct full mockup using copper wire to verify     construction steps

Schedule

1.	PDR _________________________	6/12/07

2.	FDR _________________________	8/9/07

3.	Purchase equipment for winding

mandrel  _____________________	Done

4.	Order 0.032 SS coax ____________	Received

5.     PPPL sets up winding equipment __	7/16 to 8/31/07

6.	PPPL winds and assembles 3 coils_	9/10 to 9/18/07

7.	Install on vacuum vessels  ________	 9/11 to 9/19/07

8.	Route twisted leads to JB _________    9/24 to 9/25/07

9.	Terminate leads ________________	 Oct-Nov 2007

Cost

1.	Winding equipment ___________________	$ 3000

2.	Labor to set up work area

1 techs ( 60 mh ) ________________	$ 5400

3.	Labor to wind 3 coils

3 techs( 78 mh )   _________________      $ 8000

Larry Guttadora, ( Need to ID two additional )

4.	Material Cost ___________________________       $ 4100

Cable _____	$  2000

Mandrel ___	$  1000

Nextel _____     $    450

Flex Sheath _	$    150