Certificate of Calibration

Calibration Information

Calibration Date:					14 April 2008

Target Plate Name:				PCP-TCF-007

Plate material:					Aluminum

Calibration Temperature:			68°F

Coefficient of thermal expansion		1.244 x 10-5

Target Checking Fixture

Target Capacity					50

Target Specification				0.25” shank dia.

0.500” offset

0.3125” offset

Certified by:									  	        _____________________________

Senior Product specialist

Target Layout – Target Checking Fixture – 0.500 inch offset

Point Listing – Target Checking Fixture – 0.500 inch offset

All points in millimeters

| Point Label | X | Y | Z |
| --- | --- | --- | --- |
| B1 | 151.3539 | 452.2391 | 16.19515 |
| B2 | 202.1904 | 452.0126 | 16.28274 |
| B3 | 252.9554 | 451.9295 | 16.38676 |
| B4 | 49.55658 | 401.713 | 16.43184 |
| B5 | 100.4135 | 401.5386 | 16.2891 |
| B6 | 151.227 | 401.4039 | 16.63865 |
| B7 | 202.0263 | 401.2142 | 16.72654 |
| B8 | 252.8383 | 401.0715 | 16.65335 |
| B9 | 303.6417 | 400.9312 | 16.96799 |
| B10 | 354.4378 | 400.8247 | 17.01111 |
| B11 | 49.45591 | 350.886 | 16.65567 |
| B12 | 100.3041 | 350.6617 | 16.60646 |
| B13 | 151.081 | 350.5113 | 16.84505 |
| B14 | 201.8647 | 350.3607 | 16.89229 |
| B15 | 252.6587 | 350.2155 | 17.05524 |
| B16 | 303.4414 | 350.0967 | 17.19353 |
| B17 | 354.3012 | 349.9608 | 17.35216 |
| B18 | 49.29207 | 299.9945 | 16.89157 |
| B19 | 100.1303 | 299.8003 | 16.90524 |
| B20 | 150.9439 | 299.6902 | 16.94882 |
| B21 | 201.7521 | 299.5707 | 17.00755 |
| B22 | 252.5552 | 299.4026 | 17.14626 |
| B23 | 303.3341 | 299.2177 | 17.2253 |
| B24 | 354.1281 | 299.1205 | 17.43821 |
| B25 | 49.20493 | 249.1307 | 16.92638 |
| B26 | 99.99278 | 248.9609 | 16.90907 |
| B27 | 150.7956 | 248.8167 | 16.89839 |
| B28 | 201.6014 | 248.6682 | 17.14598 |
| B29 | 252.3835 | 248.5043 | 17.40821 |
| B30 | 303.2185 | 248.3971 | 17.35721 |
| B31 | 354.0086 | 248.2279 | 17.4183 |
| B32 | 49.08307 | 198.287 | 16.95192 |
| B33 | 99.85744 | 198.133 | 16.8854 |
| B34 | 150.683 | 197.9789 | 17.07756 |
| B35 | 201.4395 | 197.8375 | 16.90117 |
| B36 | 252.2312 | 197.6393 | 17.14488 |
| B37 | 303.0431 | 197.5098 | 17.17414 |
| B38 | 353.8429 | 197.3678 | 17.35799 |
| B39 | 48.91402 | 147.4217 | 16.64878 |
| B40 | 99.71602 | 147.2175 | 16.67077 |
| B41 | 150.4882 | 147.1185 | 16.62498 |
| B42 | 201.2889 | 146.9431 | 16.83984 |
| B43 | 252.058 | 146.7558 | 16.91705 |
| B44 | 302.8999 | 146.676 | 16.89513 |
| B45 | 353.683 | 146.512 | 17.00096 |
| B46 | 99.54506 | 96.39884 | 16.29811 |
| B47 | 150.3289 | 96.24138 | 16.32538 |
| B48 | 201.1332 | 96.08101 | 16.46338 |
| B49 | 251.9492 | 95.96924 | 16.57365 |
| B50 | 302.7427 | 95.78616 | 16.66129 |
| CODE121 | 0.345805 | 553.3699 | -0.25223 |
| CODE122 | 0.247875 | 443.7951 | 0.525827 |
| CODE123 | -0.39373 | 110.9855 | 0.919354 |
| CODE124 | -1.07312 | 0.334329 | -0.32143 |
| CODE125 | 198.083 | 531.6326 | 0.387517 |
| CODE126 | 196.9655 | 19.99022 | 0.309985 |
| CODE127 | 401.1518 | 550.9183 | 0.699667 |
| CODE128 | 400.8364 | 441.6114 | 1.427955 |
| CODE129 | 399.5569 | 109.8683 | 1.565997 |
| CODE130 | 398.9883 | -0.51328 | 0.466864 |
| S1 | 72.22908 | 557.4651 | -0.0946 |
| S2 | 123.6138 | 557.5008 | 0.032572 |
| S3 | 281.7184 | 557.2375 | 0.413176 |
| S4 | 333.3555 | 557.1107 | 0.495381 |
| S9 | -11.9033 | 380.2084 | 0.893555 |
| S10 | 415.7192 | 372.6211 | 1.870929 |
| S11 | -11.8829 | 330.1545 | 1.169911 |
| S12 | 415.4893 | 321.9897 | 2.098943 |
| S13 | -11.7651 | 278.6654 | 1.343226 |
| S14 | 414.786 | 270.0804 | 2.199085 |
| S15 | -12.1569 | 228.4666 | 1.374126 |
| S16 | 415.7012 | 220.2007 | 2.155009 |
| S17 | -12.7843 | 177.0992 | 1.27745 |
| S18 | 415.0155 | 169.5171 | 1.957911 |
| S23 | 70.12703 | -8.77536 | -0.31512 |
| S24 | 121.4302 | -9.27381 | -0.24037 |
| S25 | 280.1538 | -9.24558 | 0.082716 |
| S26 | 330.9328 | -9.00329 | 0.204789 |

Target Layout – Target Checking Fixture – 0.3125 inch offset

Point Listing – Target Checking Fixture – 0.3125 inch offset

All points in millimeters

| Point Label | X | Y | Z |
| --- | --- | --- | --- |
| B1 | 151.3539 | 452.2391 | 8.257645 |
| B2 | 202.1904 | 452.0126 | 8.345235 |
| B3 | 252.9554 | 451.9295 | 8.449258 |
| B4 | 49.55658 | 401.713 | 8.494343 |
| B5 | 100.4135 | 401.5386 | 8.351598 |
| B6 | 151.227 | 401.4039 | 8.701151 |
| B7 | 202.0263 | 401.2142 | 8.789041 |
| B8 | 252.8383 | 401.0715 | 8.715854 |
| B9 | 303.6417 | 400.9312 | 9.030494 |
| B10 | 354.4378 | 400.8247 | 9.073607 |
| B11 | 49.45591 | 350.886 | 8.718175 |
| B12 | 100.3041 | 350.6617 | 8.66896 |
| B13 | 151.081 | 350.5113 | 8.907554 |
| B14 | 201.8647 | 350.3607 | 8.954791 |
| B15 | 252.6587 | 350.2155 | 9.117737 |
| B16 | 303.4414 | 350.0967 | 9.256027 |
| B17 | 354.3012 | 349.9608 | 9.414659 |
| B18 | 49.29207 | 299.9945 | 8.954067 |
| B19 | 100.1303 | 299.8003 | 8.967737 |
| B20 | 150.9439 | 299.6902 | 9.011315 |
| B21 | 201.7521 | 299.5707 | 9.070048 |
| B22 | 252.5552 | 299.4026 | 9.208758 |
| B23 | 303.3341 | 299.2177 | 9.287805 |
| B24 | 354.1281 | 299.1205 | 9.500715 |
| B25 | 49.20493 | 249.1307 | 8.988877 |
| B26 | 99.99278 | 248.9609 | 8.971571 |
| B27 | 150.7956 | 248.8167 | 8.960893 |
| B28 | 201.6014 | 248.6682 | 9.208484 |
| B29 | 252.3835 | 248.5043 | 9.47071 |
| B30 | 303.2185 | 248.3971 | 9.419707 |
| B31 | 354.0086 | 248.2279 | 9.480797 |
| B32 | 49.08307 | 198.287 | 9.014423 |
| B33 | 99.85744 | 198.133 | 8.947896 |
| B34 | 150.683 | 197.9789 | 9.140063 |
| B35 | 201.4395 | 197.8375 | 8.963666 |
| B36 | 252.2312 | 197.6393 | 9.207377 |
| B37 | 303.0431 | 197.5098 | 9.236642 |
| B38 | 353.8429 | 197.3678 | 9.420488 |
| B39 | 48.91402 | 147.4217 | 8.711276 |
| B40 | 99.71602 | 147.2175 | 8.733265 |
| B41 | 150.4882 | 147.1185 | 8.68748 |
| B42 | 201.2889 | 146.9431 | 8.902339 |
| B43 | 252.058 | 146.7558 | 8.97955 |
| B44 | 302.8999 | 146.676 | 8.957628 |
| B45 | 353.683 | 146.512 | 9.063458 |
| B46 | 99.54506 | 96.39884 | 8.360608 |
| B47 | 150.3289 | 96.24138 | 8.387877 |
| B48 | 201.1332 | 96.08101 | 8.525882 |
| B49 | 251.9492 | 95.96924 | 8.636149 |
| B50 | 302.7427 | 95.78616 | 8.72379 |
| CODE121 | 0.345805 | 553.3699 | -0.25223 |
| CODE122 | 0.247875 | 443.7951 | 0.525827 |
| CODE123 | -0.39373 | 110.9855 | 0.919354 |
| CODE124 | -1.07312 | 0.334329 | -0.32143 |
| CODE125 | 198.083 | 531.6326 | 0.387517 |
| CODE126 | 196.9655 | 19.99022 | 0.309985 |
| CODE127 | 401.1518 | 550.9183 | 0.699667 |
| CODE128 | 400.8364 | 441.6114 | 1.427955 |
| CODE129 | 399.5569 | 109.8683 | 1.565997 |
| CODE130 | 398.9883 | -0.51328 | 0.466864 |
| S1 | 72.22908 | 557.4651 | -0.0946 |
| S2 | 123.6138 | 557.5008 | 0.032572 |
| S3 | 281.7184 | 557.2375 | 0.413176 |
| S4 | 333.3555 | 557.1107 | 0.495381 |
| S9 | -11.9033 | 380.2084 | 0.893555 |
| S10 | 415.7192 | 372.6211 | 1.870929 |
| S11 | -11.8829 | 330.1545 | 1.169911 |
| S12 | 415.4893 | 321.9897 | 2.098943 |
| S13 | -11.7651 | 278.6654 | 1.343226 |
| S14 | 414.786 | 270.0804 | 2.199085 |
| S15 | -12.1569 | 228.4666 | 1.374126 |
| S16 | 415.7012 | 220.2007 | 2.155009 |
| S17 | -12.7843 | 177.0992 | 1.27745 |
| S18 | 415.0155 | 169.5171 | 1.957911 |
| S23 | 70.12703 | -8.77536 | -0.31512 |
| S24 | 121.4302 | -9.27381 | -0.24037 |
| S25 | 280.1538 | -9.24558 | 0.082716 |
| S26 | 330.9328 | -9.00329 | 0.204789 |

# Text boxes

Y

X

Y

X