## Master part
2007-07-12 20:25
P.O. Box 451
Princeton, NJ 08543
(609)243-3626
WWW.PPPL.GOV
INSPECTION REPORT
PART NO. | [5 empty] | CUSTOMER
REV. NO. | [5 empty] | CONTACT NAME
PART NAME | [5 empty] | CONTACT PHONE
PROD. DATE | [5 empty] | PURCH. ORDER#
DATE REC'D | [5 empty] | TEMPERATURE
QTY. REC'D | [5 empty] | ARM SERIAL #
ACCEPTED
REJECTED
EMAIL TO: | [5 empty] | FAX TO:
INSPECTED BY:
REMARKS AND DISPOSITION
Conical Seats (averaged-best fit)
Conical Seat 2
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 76.7512 |  | 76.7491 |  | -0.0020641
Target point |  | Y | -0.01 | 0.01 | -37.0189 |  | -37.0176 |  | 0.0013173
|  | Z | -0.01 | 0.01 | -25.0018 |  | -25.0099 |  | -0.0081052
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.008467 |  | 0.008467
Conical Seat 3
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 96.1272 |  | 96.1212 |  | -0.0059977
Target point |  | Y | -0.01 | 0.01 | -15.6549 |  | -15.6617 |  | -0.0067504
|  | Z | -0.01 | 0.01 | -24.4919 |  | -24.4883 |  | 0.0035631
Distance To Target Point |  |  | -0.005 | 0.005 | 0 |  | 0.0097075 |  | 0.0097075 |  | 0.0047075
Conical Seat 5
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 79.7088 |  | 79.7073 |  | -0.001522
Target point |  | Y | -0.01 | 0.01 | 30.9411 |  | 30.9433 |  | 0.0021939
|  | Z | -0.01 | 0.01 | -27.5135 |  | -27.5078 |  | 0.0056932
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0062883 |  | 0.0062883
Conical Seat 7
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 58.2322 |  | 58.2322 |  | 2.825e-05
Target point |  | Y | -0.01 | 0.01 | 48.5937 |  | 48.5973 |  | 0.0036188
|  | Z | -0.01 | 0.01 | -9.7523 |  | -9.7498 |  | 0.0025134
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0044061 |  | 0.0044061
Conical Seat 8
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 41.6065 |  | 41.6053 |  | -0.001188
Target point |  | Y | -0.01 | 0.01 | 49.9905 |  | 49.995 |  | 0.0044718
|  | Z | -0.01 | 0.01 | -6.2829 |  | -6.2835 |  | -0.00061104
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0046671 |  | 0.0046671
Conical Seat 9
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 21.3958 |  | 21.3991 |  | 0.0032725
Target point |  | Y | -0.01 | 0.01 | 29.1604 |  | 29.1585 |  | -0.0018725
|  | Z | -0.01 | 0.01 | -3.7846 |  | -3.7909 |  | -0.0062867
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0073307 |  | 0.0073307
Conical Seat 10
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 35.2895 |  | 35.2894 |  | -0.00014895
Target point |  | Y | -0.01 | 0.01 | -2.2294 |  | -2.2302 |  | -0.00082603
|  | Z | -0.01 | 0.01 | -8.9735 |  | -8.9791 |  | -0.00562
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0056823 |  | 0.0056823
Conical Seat 11
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 23.1781 |  | 23.1829 |  | 0.0047596
Target point |  | Y | -0.01 | 0.01 | -25.0226 |  | -25.0253 |  | -0.0027113
|  | Z | -0.01 | 0.01 | -2.602 |  | -2.5975 |  | 0.0044652
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.007067 |  | 0.007067
Conical Seat 12
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 34.0791 |  | 34.082 |  | 0.0028604
Target point |  | Y | -0.01 | 0.01 | -45.8046 |  | -45.804 |  | 0.00055839
|  | Z | -0.01 | 0.01 | -3.0653 |  | -3.0609 |  | 0.0043881
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0052677 |  | 0.0052677
Inspection Group 33
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 54.1817 | 40.8243 | -15.4819 | 0.007 | 0.0406 | -0.0184 | -0.0452
| SP-14 | 0 | -0.01 | 0.01 | 54.1387 | 40.6539 | -15.8755 | 0.0052 | 0.03 | -0.0136 | -0.0334
| SP-15 | 0 | -0.01 | 0.01 | 54.1084 | 40.5204 | -16.1827 | 0.0036 | 0.0212 | -0.0096 | -0.0235
| SP-16 | 0 | -0.01 | 0.01 | 54.0877 | 40.3746 | -16.5133 | -0.0003 | -0.0019 | 0.0008 | 0.0021
| SP-17 | 0 | -0.01 | 0.01 | 53.2405 | 41.0121 | -15.4009 | 0.0057 | 0.0382 | -0.0172 | -0.0423
| SP-18 | 0 | -0.01 | 0.01 | 53.1918 | 40.848 | -15.7807 | 0.0039 | 0.0268 | -0.0121 | -0.0296
| SP-19 | 0 | -0.01 | 0.01 | 53.1471 | 40.7117 | -16.0978 | 0.0041 | 0.0277 | -0.0125 | -0.0307
| SP-20 | 0 | -0.01 | 0.01 | 53.1386 | 40.5371 | -16.4881 | -0.0003 | -0.0021 | 0.0009 | 0.0023
| SP-21 | 0 | -0.01 | 0.01 | 52.5145 | 41.1227 | -15.3792 | 0.0073 | 0.0564 | -0.0254 | -0.0623
| SP-22 | 0 | -0.01 | 0.01 | 52.4634 | 40.9892 | -15.6903 | 0.0062 | 0.0481 | -0.0217 | -0.0531
| SP-23 | 0 | -0.01 | 0.01 | 52.4044 | 40.8408 | -16.0369 | 0.0051 | 0.0404 | -0.0182 | -0.0446
| SP-24 | 0 | -0.01 | 0.01 | 52.3292 | 40.6899 | -16.3933 | 0.0022 | 0.0172 | -0.0077 | -0.019
Number of points |  |  | 12
Out of Tolerance |  |  | 10
Performance |  |  | 16.67%
| Mean |  | -0.031608
| Std. deviation |  | 0.019978
| Max. Value |  | 0.0023
| Min. Value |  | -0.0623
Inspection Group 34
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 51.55 | 41.2907 | -15.2544 | 0.0071 | 0.0698 | -0.0314 | -0.0768
| SP-14 | 0 | -0.01 | 0.01 | 51.4914 | 41.1433 | -15.5957 | 0.0063 | 0.0626 | -0.0281 | -0.0689
| SP-15 | 0 | -0.01 | 0.01 | 51.4395 | 41.0101 | -15.9039 | 0.0045 | 0.0448 | -0.0201 | -0.0493
| SP-16 | 0 | -0.01 | 0.01 | 51.3962 | 40.8561 | -16.2566 | 0.0032 | 0.032 | -0.0143 | -0.0352
| SP-17 | 0 | -0.01 | 0.01 | 50.902 | 41.4104 | -15.1219 | 0.0054 | 0.0646 | -0.029 | -0.071
| SP-18 | 0 | -0.01 | 0.01 | 50.8261 | 41.2552 | -15.4817 | 0.0045 | 0.0547 | -0.0245 | -0.0601
| SP-19 | 0 | -0.01 | 0.01 | 50.7267 | 41.1221 | -15.7967 | 0.002 | 0.0258 | -0.0116 | -0.0283
| SP-20 | 0 | -0.01 | 0.01 | 50.6085 | 40.9707 | -16.1551 | 0.0015 | 0.0202 | -0.009 | -0.0222
| SP-21 | 0 | -0.01 | 0.01 | 49.9717 | 41.5508 | -14.9519 | 0.0043 | 0.0794 | -0.0355 | -0.0871
| SP-22 | 0 | -0.01 | 0.01 | 49.8694 | 41.4192 | -15.2583 | 0.0037 | 0.0706 | -0.0316 | -0.0775
| SP-23 | 0 | -0.01 | 0.01 | 49.7716 | 41.2861 | -15.5674 | 0.0025 | 0.0515 | -0.023 | -0.0564
| SP-24 | 0 | -0.01 | 0.01 | 49.6842 | 41.1646 | -15.8493 | 0.0017 | 0.0364 | -0.0162 | -0.0399
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.056058
| Std. deviation |  | 0.021142
| Max. Value |  | -0.0222
| Min. Value |  | -0.0871
Inspection Group 35
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 49.0411 | 41.7154 | -14.6637 | 0.0018 | 0.0787 | -0.035 | -0.0862
| SP-14 | 0 | -0.01 | 0.01 | 48.9348 | 41.572 | -14.9913 | 0.0014 | 0.0711 | -0.0316 | -0.0778
| SP-15 | 0 | -0.01 | 0.01 | 48.8274 | 41.4262 | -15.3246 | 0.0011 | 0.063 | -0.0279 | -0.0689
| SP-16 | 0 | -0.01 | 0.01 | 48.7588 | 41.2765 | -15.665 | 0.0007 | 0.0399 | -0.0177 | -0.0437
| SP-17 | 0 | -0.01 | 0.01 | 48.2042 | 41.8578 | -14.3568 | -0.0006 | 0.0735 | -0.0324 | -0.0803
| SP-18 | 0 | -0.01 | 0.01 | 48.0434 | 41.7142 | -14.6791 | -0.0008 | 0.0715 | -0.0315 | -0.0782
| SP-19 | 0 | -0.01 | 0.01 | 47.9352 | 41.5681 | -15.0081 | -0.0009 | 0.0668 | -0.0294 | -0.073
| SP-20 | 0 | -0.01 | 0.01 | 47.819 | 41.4059 | -15.3737 | -0.0007 | 0.0472 | -0.0207 | -0.0515
| SP-21 | 0 | -0.01 | 0.01 | 47.5318 | 41.9369 | -14.1461 | -0.0032 | 0.0986 | -0.0432 | -0.1077
| SP-22 | 0 | -0.01 | 0.01 | 47.3824 | 41.8104 | -14.4237 | -0.0036 | 0.0997 | -0.0435 | -0.1089
| SP-23 | 0 | -0.01 | 0.01 | 47.2232 | 41.6772 | -14.7152 | -0.0038 | 0.096 | -0.0418 | -0.1048
| SP-24 | 0 | -0.01 | 0.01 | 47.0317 | 41.5235 | -15.0504 | -0.0031 | 0.0706 | -0.0306 | -0.077
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.079833
| Std. deviation |  | 0.020394
| Max. Value |  | -0.0437
| Min. Value |  | -0.1089
Inspection Group 36
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 46.6374 | 42.0622 | -13.7549 | -0.0074 | 0.1097 | -0.0474 | -0.1197
| SP-14 | 0 | -0.01 | 0.01 | 46.4661 | 41.9186 | -14.0603 | -0.0071 | 0.0994 | -0.0428 | -0.1085
| SP-15 | 0 | -0.01 | 0.01 | 46.3238 | 41.7965 | -14.3203 | -0.0068 | 0.0906 | -0.0389 | -0.0988
| SP-16 | 0 | -0.01 | 0.01 | 46.1712 | 41.662 | -14.6073 | -0.0055 | 0.0704 | -0.0301 | -0.0768
| SP-17 | 0 | -0.01 | 0.01 | 45.8616 | 42.1859 | -13.315 | -0.0085 | 0.0849 | -0.0362 | -0.0927
| SP-18 | 0 | -0.01 | 0.01 | 45.7379 | 42.0408 | -13.627 | -0.008 | 0.0783 | -0.0332 | -0.0854
| SP-19 | 0 | -0.01 | 0.01 | 45.5747 | 41.8866 | -13.9512 | -0.0053 | 0.0505 | -0.0214 | -0.0551
| SP-20 | 0 | -0.01 | 0.01 | 45.461 | 41.7445 | -14.2593 | -0.0029 | 0.0278 | -0.0117 | -0.0303
| SP-21 | 0 | -0.01 | 0.01 | 45.0592 | 42.2344 | -12.9798 | -0.0127 | 0.096 | -0.0402 | -0.1049
| SP-22 | 0 | -0.01 | 0.01 | 44.9232 | 42.1121 | -13.2288 | -0.0103 | 0.0762 | -0.0318 | -0.0832
| SP-23 | 0 | -0.01 | 0.01 | 44.7278 | 41.9382 | -13.5823 | -0.0069 | 0.0496 | -0.0206 | -0.0542
| SP-24 | 0 | -0.01 | 0.01 | 44.5755 | 41.8096 | -13.841 | -0.0022 | 0.0152 | -0.0063 | -0.0166
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.077183
| Std. deviation |  | 0.031984
| Max. Value |  | -0.0166
| Min. Value |  | -0.1197
Inspection Group 37
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 44.2335 | 42.3301 | -12.451 | -0.013 | 0.0777 | -0.0319 | -0.085
| SP-14 | 0 | -0.01 | 0.01 | 44.0547 | 42.1716 | -12.7638 | -0.0106 | 0.0622 | -0.0255 | -0.0681
| SP-15 | 0 | -0.01 | 0.01 | 43.898 | 42.0291 | -13.0464 | -0.0069 | 0.0394 | -0.0161 | -0.0431
| SP-16 | 0 | -0.01 | 0.01 | 43.7343 | 41.8688 | -13.3699 | -0.002 | 0.0114 | -0.0046 | -0.0125
| SP-17 | 0 | -0.01 | 0.01 | 43.4058 | 42.3359 | -12.0616 | -0.0089 | 0.044 | -0.0177 | -0.0482
| SP-18 | 0 | -0.01 | 0.01 | 43.1941 | 42.1964 | -12.3011 | -0.006 | 0.0291 | -0.0117 | -0.0319
| SP-19 | 0 | -0.01 | 0.01 | 43.0271 | 42.0611 | -12.5519 | -0.0044 | 0.021 | -0.0084 | -0.0231
| SP-20 | 0 | -0.01 | 0.01 | 42.9007 | 41.9023 | -12.8831 | 0.0021 | -0.0098 | 0.0039 | 0.0107
| SP-21 | 0 | -0.01 | 0.01 | 42.6437 | 42.3398 | -11.6364 | -0.0107 | 0.0455 | -0.018 | -0.0501
| SP-22 | 0 | -0.01 | 0.01 | 42.4802 | 42.1824 | -11.9372 | -0.0087 | 0.0366 | -0.0144 | -0.0403
| SP-23 | 0 | -0.01 | 0.01 | 42.3193 | 42.0301 | -12.2264 | -0.0051 | 0.0211 | -0.0083 | -0.0233
| SP-24 | 0 | -0.01 | 0.01 | 42.1896 | 41.9125 | -12.4463 | -0.001 | 0.004 | -0.0016 | -0.0044
Number of points |  |  | 12
Out of Tolerance |  |  | 11
Performance |  |  | 8.333%
| Mean |  | -0.034942
| Std. deviation |  | 0.026817
| Max. Value |  | 0.0107
| Min. Value |  | -0.085
Inspection Group 38
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 41.8132 | 42.3362 | -11.11 | -0.0116 | 0.0426 | -0.0166 | -0.0472
| SP-14 | 0 | -0.01 | 0.01 | 41.6475 | 42.1774 | -11.4028 | -0.0091 | 0.0332 | -0.0128 | -0.0367
| SP-15 | 0 | -0.01 | 0.01 | 41.4892 | 42.0249 | -11.6843 | -0.0052 | 0.0188 | -0.0072 | -0.0208
| SP-16 | 0 | -0.01 | 0.01 | 41.3369 | 41.8624 | -11.9962 | -0.0003 | 0.0012 | -0.0005 | -0.0014
| SP-17 | 0 | -0.01 | 0.01 | 41.0254 | 42.2781 | -10.6668 | -0.0135 | 0.0436 | -0.0166 | -0.0486
| SP-18 | 0 | -0.01 | 0.01 | 40.8523 | 42.136 | -10.8987 | -0.0118 | 0.0375 | -0.0142 | -0.0418
| SP-19 | 0 | -0.01 | 0.01 | 40.6714 | 41.9558 | -11.2237 | -0.0027 | 0.0084 | -0.0032 | -0.0094
| SP-20 | 0 | -0.01 | 0.01 | 40.597 | 41.8191 | -11.5239 | 0.0043 | -0.0135 | 0.0051 | 0.015
| SP-21 | 0 | -0.01 | 0.01 | 40.203 | 42.2158 | -10.112 | -0.0263 | 0.0747 | -0.0279 | -0.084
| SP-22 | 0 | -0.01 | 0.01 | 40.0476 | 42.0394 | -10.4379 | -0.0215 | 0.0604 | -0.0224 | -0.0679
| SP-23 | 0 | -0.01 | 0.01 | 39.9216 | 41.9005 | -10.691 | -0.0129 | 0.0358 | -0.0132 | -0.0403
| SP-24 | 0 | -0.01 | 0.01 | 39.7699 | 41.735 | -10.9905 | -0.0029 | 0.0078 | -0.0029 | -0.0088
Number of points |  |  | 12
Out of Tolerance |  |  | 9
Performance |  |  | 25%
| Mean |  | -0.032658
| Std. deviation |  | 0.028606
| Max. Value |  | 0.015
| Min. Value |  | -0.084
Inspection Group 39
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-15 | 0 | -0.01 | 0.01 | 39.3483 | 42.0574 | -9.6688 | -0.0339 | 0.0844 | -0.0307 | -0.096
| SP-16 | 0 | -0.01 | 0.01 | 39.2286 | 41.9226 | -9.9067 | -0.0295 | 0.0727 | -0.0263 | -0.0827
| SP-17 | 0 | -0.01 | 0.01 | 39.0487 | 41.7193 | -10.2659 | -0.0227 | 0.0551 | -0.0198 | -0.0628
| SP-18 | 0 | -0.01 | 0.01 | 38.9311 | 41.587 | -10.4987 | -0.0124 | 0.0298 | -0.0107 | -0.034
| SP-19 | 0 | -0.01 | 0.01 | 38.7657 | 41.9739 | -9.2183 | -0.0316 | 0.0715 | -0.0255 | -0.0822
| SP-20 | 0 | -0.01 | 0.01 | 38.6081 | 41.8006 | -9.5085 | -0.0318 | 0.071 | -0.0251 | -0.0817
| SP-21 | 0 | -0.01 | 0.01 | 38.456 | 41.625 | -9.8113 | -0.0272 | 0.06 | -0.0211 | -0.0692
| SP-22 | 0 | -0.01 | 0.01 | 38.2839 | 41.4459 | -10.097 | -0.0129 | 0.028 | -0.0098 | -0.0324
| SP-23 | 0 | -0.01 | 0.01 | 37.9681 | 41.7803 | -8.6968 | -0.054 | 0.1072 | -0.037 | -0.1256
| SP-24 | 0 | -0.01 | 0.01 | 37.8178 | 41.5834 | -9.0484 | -0.0489 | 0.096 | -0.0329 | -0.1126
| SP-25 | 0 | -0.01 | 0.01 | 37.7136 | 41.4499 | -9.283 | -0.0477 | 0.0929 | -0.0317 | -0.1091
| SP-26 | 0 | -0.01 | 0.01 | 37.5697 | 41.274 | -9.5822 | -0.0299 | 0.0577 | -0.0195 | -0.0679
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.079683
| Std. deviation |  | 0.028919
| Max. Value |  | -0.0324
| Min. Value |  | -0.1256
Inspection Group 40
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-25 | 0 | -0.01 | 0.01 | 37.1936 | 41.5496 | -8.141 | -0.0575 | 0.0987 | -0.0325 | -0.1188
| SP-26 | 0 | -0.01 | 0.01 | 37.0562 | 41.3695 | -8.4451 | -0.0564 | 0.0958 | -0.0313 | -0.1155
| SP-27 | 0 | -0.01 | 0.01 | 36.9124 | 41.1775 | -8.7737 | -0.0507 | 0.0851 | -0.0276 | -0.1028
| SP-28 | 0 | -0.01 | 0.01 | 36.7919 | 40.9932 | -9.1216 | -0.0345 | 0.0575 | -0.0185 | -0.0695
| SP-29 | 0 | -0.01 | 0.01 | 36.6158 | 41.3346 | -7.6887 | -0.0542 | 0.0812 | -0.025 | -0.1008
| SP-30 | 0 | -0.01 | 0.01 | 36.5163 | 41.1631 | -8.0304 | -0.0519 | 0.0778 | -0.0239 | -0.0965
| SP-31 | 0 | -0.01 | 0.01 | 36.3574 | 40.9738 | -8.3001 | -0.0384 | 0.0565 | -0.0171 | -0.0705
| SP-32 | 0 | -0.01 | 0.01 | 36.2326 | 40.7848 | -8.6446 | -0.0249 | 0.0365 | -0.011 | -0.0456
| SP-33 | 0 | -0.01 | 0.01 | 36.1247 | 41.1072 | -7.2678 | -0.0695 | 0.0904 | -0.0251 | -0.1167
| SP-34 | 0 | -0.01 | 0.01 | 35.9813 | 40.9142 | -7.5653 | -0.068 | 0.0875 | -0.024 | -0.1134
| SP-35 | 0 | -0.01 | 0.01 | 35.8231 | 40.7052 | -7.8789 | -0.0525 | 0.0668 | -0.018 | -0.0868
| SP-36 | 0 | -0.01 | 0.01 | 35.6577 | 40.4977 | -8.1639 | -0.0324 | 0.0405 | -0.0107 | -0.0529
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.090817
| Std. deviation |  | 0.025608
| Max. Value |  | -0.0456
| Min. Value |  | -0.1188
Inspection Group 41
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 35.4826 | 40.7146 | -6.6349 | -0.048 | 0.0505 | -0.0108 | -0.0705
| SP-14 | 0 | -0.01 | 0.01 | 35.3389 | 40.5187 | -6.9134 | -0.0542 | 0.0566 | -0.0119 | -0.0792
| SP-15 | 0 | -0.01 | 0.01 | 35.1854 | 40.3052 | -7.231 | -0.0469 | 0.0487 | -0.0101 | -0.0684
| SP-16 | 0 | -0.01 | 0.01 | 35.0413 | 40.1045 | -7.5311 | -0.0308 | 0.0318 | -0.0065 | -0.0448
| SP-17 | 0 | -0.01 | 0.01 | 35.1689 | 40.4653 | -6.2932 | -0.0328 | 0.0309 | -0.0052 | -0.0453
| SP-18 | 0 | -0.01 | 0.01 | 35.0473 | 40.3047 | -6.4799 | -0.0381 | 0.0355 | -0.0058 | -0.0524
| SP-19 | 0 | -0.01 | 0.01 | 34.8817 | 40.0815 | -6.7588 | -0.0308 | 0.0284 | -0.0045 | -0.0421
| SP-20 | 0 | -0.01 | 0.01 | 34.705 | 39.8419 | -7.0644 | -0.0106 | 0.0097 | -0.0015 | -0.0145
| SP-21 | 0 | -0.01 | 0.01 | 34.9071 | 40.2317 | -5.8557 | -0.0386 | 0.0323 | -0.0035 | -0.0505
| SP-22 | 0 | -0.01 | 0.01 | 34.7695 | 40.0439 | -6.0688 | -0.0435 | 0.036 | -0.0037 | -0.0566
| SP-23 | 0 | -0.01 | 0.01 | 34.5718 | 39.7746 | -6.3639 | -0.0401 | 0.0327 | -0.003 | -0.0518
| SP-24 | 0 | -0.01 | 0.01 | 34.4122 | 39.5576 | -6.5932 | -0.0257 | 0.0207 | -0.0018 | -0.033
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.050758
| Std. deviation |  | 0.017368
| Max. Value |  | -0.0145
| Min. Value |  | -0.0792
Inspection Group 42
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 34.4828 | 39.7127 | -5.1296 | -0.0605 | 0.0427 | -0.0002 | -0.0741
| SP-14 | 0 | -0.01 | 0.01 | 34.3392 | 39.5084 | -5.3363 | -0.0583 | 0.0409 | 0 | -0.0712
| SP-15 | 0 | -0.01 | 0.01 | 34.1581 | 39.2505 | -5.6001 | -0.056 | 0.0389 | 0.0004 | -0.0682
| SP-16 | 0 | -0.01 | 0.01 | 34.0189 | 39.052 | -5.8016 | -0.0486 | 0.0336 | 0.0005 | -0.0591
| SP-17 | 0 | -0.01 | 0.01 | 34.2078 | 39.2892 | -4.3418 | -0.0617 | 0.0401 | 0.0016 | -0.0736
| SP-18 | 0 | -0.01 | 0.01 | 34.0257 | 39.0198 | -4.6064 | -0.0635 | 0.0409 | 0.002 | -0.0756
| SP-19 | 0 | -0.01 | 0.01 | 33.8457 | 38.7516 | -4.8395 | -0.0517 | 0.033 | 0.0019 | -0.0614
| SP-20 | 0 | -0.01 | 0.01 | 33.7107 | 38.5554 | -5.0971 | -0.0348 | 0.0221 | 0.0014 | -0.0413
| SP-21 | 0 | -0.01 | 0.01 | 33.9282 | 38.8189 | -3.734 | -0.0784 | 0.0488 | 0.0027 | -0.0924
| SP-22 | 0 | -0.01 | 0.01 | 33.7642 | 38.5687 | -3.9813 | -0.0756 | 0.0467 | 0.0029 | -0.0889
| SP-23 | 0 | -0.01 | 0.01 | 33.6101 | 38.3336 | -4.2121 | -0.0664 | 0.0408 | 0.0028 | -0.078
| SP-24 | 0 | -0.01 | 0.01 | 33.4621 | 38.1072 | -4.4225 | -0.0436 | 0.0266 | 0.002 | -0.0512
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.069583
| Std. deviation |  | 0.014637
| Max. Value |  | -0.0413
| Min. Value |  | -0.0924
Inspection Group 43
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 33.6357 | 38.2908 | -2.9548 | -0.0663 | 0.0395 | 0.0027 | -0.0773
| SP-14 | 0 | -0.01 | 0.01 | 33.4829 | 38.0499 | -3.1892 | -0.0563 | 0.0333 | 0.0025 | -0.0654
| SP-15 | 0 | -0.01 | 0.01 | 33.3445 | 37.8321 | -3.4039 | -0.0542 | 0.0319 | 0.0025 | -0.063
| SP-16 | 0 | -0.01 | 0.01 | 33.1691 | 37.5577 | -3.6945 | -0.0312 | 0.0183 | 0.0016 | -0.0362
| SP-17 | 0 | -0.01 | 0.01 | 33.3863 | 37.8081 | -2.1776 | -0.0403 | 0.0232 | 0.0018 | -0.0466
| SP-18 | 0 | -0.01 | 0.01 | 33.2044 | 37.5141 | -2.4709 | -0.0347 | 0.0198 | 0.0016 | -0.04
| SP-19 | 0 | -0.01 | 0.01 | 33.0929 | 37.341 | -2.7337 | -0.035 | 0.0199 | 0.0018 | -0.0403
| SP-20 | 0 | -0.01 | 0.01 | 32.95 | 37.115 | -3.0109 | -0.0065 | 0.0037 | 0.0004 | -0.0075
| SP-21 | 0 | -0.01 | 0.01 | 33.0887 | 37.2231 | -1.4574 | -0.0625 | 0.0346 | 0.003 | -0.0715
| SP-22 | 0 | -0.01 | 0.01 | 32.9467 | 36.9853 | -1.6788 | -0.0489 | 0.0269 | 0.0025 | -0.0559
| SP-23 | 0 | -0.01 | 0.01 | 32.8142 | 36.7693 | -1.9465 | -0.0407 | 0.0222 | 0.0022 | -0.0464
| SP-24 | 0 | -0.01 | 0.01 | 32.7191 | 36.6205 | -2.1986 | -0.0092 | 0.005 | 0.0005 | -0.0105
Number of points |  |  | 12
Out of Tolerance |  |  | 11
Performance |  |  | 8.333%
| Mean |  | -0.046717
| Std. deviation |  | 0.021939
| Max. Value |  | -0.0075
| Min. Value |  | -0.0773
Inspection Group 44
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 32.7688 | 36.5514 | -0.6862 | -0.0674 | 0.035 | 0.0041 | -0.076
| SP-14 | 0 | -0.01 | 0.01 | 32.6633 | 36.3757 | -0.9201 | -0.0527 | 0.0272 | 0.0033 | -0.0594
| SP-15 | 0 | -0.01 | 0.01 | 32.5406 | 36.1718 | -1.1947 | -0.0381 | 0.0196 | 0.0025 | -0.043
| SP-16 | 0 | -0.01 | 0.01 | 32.4154 | 35.9664 | -1.4907 | -0.0158 | 0.0081 | 0.0011 | -0.0178
| SP-17 | 0 | -0.01 | 0.01 | 32.5521 | 36.0163 | 0.035 | -0.0711 | 0.034 | 0.0057 | -0.079
| SP-18 | 0 | -0.01 | 0.01 | 32.3998 | 35.7499 | -0.2748 | -0.0558 | 0.0264 | 0.0047 | -0.0619
| SP-19 | 0 | -0.01 | 0.01 | 32.2958 | 35.5759 | -0.5313 | -0.0243 | 0.0115 | 0.0021 | -0.027
| SP-20 | 0 | -0.01 | 0.01 | 32.1771 | 35.3783 | -0.8282 | 0.0172 | -0.0081 | -0.0015 | 0.0191
| SP-21 | 0 | -0.01 | 0.01 | 32.3038 | 35.3479 | 0.6213 | -0.112 | 0.0477 | 0.012 | -0.1224
| SP-22 | 0 | -0.01 | 0.01 | 32.1777 | 35.1247 | 0.3304 | -0.0899 | 0.0379 | 0.0099 | -0.0981
| SP-23 | 0 | -0.01 | 0.01 | 32.0629 | 34.9222 | 0.0612 | -0.0639 | 0.0267 | 0.0072 | -0.0697
| SP-24 | 0 | -0.01 | 0.01 | 31.9311 | 34.6802 | -0.2159 | -0.0067 | 0.0027 | 0.0008 | -0.0072
Number of points |  |  | 12
Out of Tolerance |  |  | 11
Performance |  |  | 8.333%
| Mean |  | -0.053533
| Std. deviation |  | 0.040184
| Max. Value |  | 0.0191
| Min. Value |  | -0.1224
Inspection Group 45
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 32.0857 | 34.5838 | 1.2601 | -0.0908 | 0.0319 | 0.013 | -0.0971
| SP-14 | 0 | -0.01 | 0.01 | 31.986 | 34.4094 | 0.9903 | -0.0897 | 0.0311 | 0.013 | -0.0959
| SP-15 | 0 | -0.01 | 0.01 | 31.8581 | 34.1701 | 0.6789 | -0.0724 | 0.0245 | 0.0108 | -0.0772
| SP-16 | 0 | -0.01 | 0.01 | 31.7422 | 33.9744 | 0.3465 | -0.0216 | 0.0072 | 0.0033 | -0.023
| SP-17 | 0 | -0.01 | 0.01 | 31.9438 | 33.8477 | 1.7951 | -0.0648 | 0.0172 | 0.0117 | -0.068
| SP-18 | 0 | -0.01 | 0.01 | 31.8353 | 33.6756 | 1.4468 | -0.0698 | 0.0184 | 0.0127 | -0.0733
| SP-19 | 0 | -0.01 | 0.01 | 31.7144 | 33.4336 | 1.1288 | -0.0475 | 0.012 | 0.0089 | -0.0498
| SP-20 | 0 | -0.01 | 0.01 | 31.623 | 33.3041 | 0.8139 | 0.0181 | -0.0046 | -0.0034 | 0.019
| SP-21 | 0 | -0.01 | 0.01 | 31.8678 | 33.17 | 2.1711 | -0.0804 | 0.0142 | 0.0173 | -0.0835
| SP-22 | 0 | -0.01 | 0.01 | 31.7722 | 32.9983 | 1.8681 | -0.0856 | 0.0146 | 0.0187 | -0.0888
| SP-23 | 0 | -0.01 | 0.01 | 31.6624 | 32.8468 | 1.4842 | -0.0675 | 0.0115 | 0.0148 | -0.0701
| SP-24 | 0 | -0.01 | 0.01 | 31.5781 | 32.7162 | 1.2005 | -0.0235 | 0.0039 | 0.0052 | -0.0243
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.061
| Std. deviation |  | 0.035242
| Max. Value |  | 0.019
| Min. Value |  | -0.0971
Inspection Group 46
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 31.8641 | 32.2341 | 2.6184 | -0.067 | 0.0038 | 0.0173 | -0.0693
| SP-14 | 0 | -0.01 | 0.01 | 31.7803 | 32.1092 | 2.3211 | -0.0806 | 0.0043 | 0.0209 | -0.0834
| SP-15 | 0 | -0.01 | 0.01 | 31.6923 | 31.9841 | 2.0079 | -0.0672 | 0.0035 | 0.0175 | -0.0696
| SP-16 | 0 | -0.01 | 0.01 | 31.5716 | 31.8249 | 1.576 | -0.0253 | 0.0013 | 0.0066 | -0.0262
| SP-17 | 0 | -0.01 | 0.01 | 31.9421 | 31.641 | 2.9377 | -0.038 | -0.001 | 0.0109 | -0.0396
| SP-18 | 0 | -0.01 | 0.01 | 31.8412 | 31.4738 | 2.5693 | -0.0491 | -0.0016 | 0.0142 | -0.0511
| SP-19 | 0 | -0.01 | 0.01 | 31.7309 | 31.3086 | 2.168 | -0.0311 | -0.0011 | 0.009 | -0.0324
| SP-20 | 0 | -0.01 | 0.01 | 31.6421 | 31.2452 | 1.8536 | 0.0031 | 0.0001 | -0.0009 | 0.0032
| SP-21 | 0 | -0.01 | 0.01 | 32.0299 | 31.0606 | 3.1028 | -0.0589 | -0.0061 | 0.0184 | -0.062
| SP-22 | 0 | -0.01 | 0.01 | 31.9614 | 30.9365 | 2.842 | -0.0552 | -0.006 | 0.0173 | -0.0582
| SP-23 | 0 | -0.01 | 0.01 | 31.8683 | 30.7793 | 2.4911 | -0.0403 | -0.0045 | 0.0127 | -0.0426
| SP-24 | 0 | -0.01 | 0.01 | 31.7886 | 30.6425 | 2.1897 | -0.0121 | -0.0014 | 0.0038 | -0.0128
Number of points |  |  | 12
Out of Tolerance |  |  | 11
Performance |  |  | 8.333%
| Mean |  | -0.045333
| Std. deviation |  | 0.025364
| Max. Value |  | 0.0032
| Min. Value |  | -0.0834
Inspection Group 47
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 32.2989 | 30.1701 | 3.4854 | -0.0481 | -0.0102 | 0.0168 | -0.052
| SP-14 | 0 | -0.01 | 0.01 | 32.2252 | 30.0427 | 3.1967 | -0.0561 | -0.012 | 0.0196 | -0.0607
| SP-15 | 0 | -0.01 | 0.01 | 32.1384 | 29.8914 | 2.8553 | -0.045 | -0.0098 | 0.0158 | -0.0487
| SP-16 | 0 | -0.01 | 0.01 | 32.0613 | 29.7669 | 2.5581 | -0.0306 | -0.0068 | 0.0108 | -0.0332
| SP-17 | 0 | -0.01 | 0.01 | 32.6097 | 29.3484 | 3.7695 | -0.008 | -0.0023 | 0.003 | -0.0088
| SP-18 | 0 | -0.01 | 0.01 | 32.5354 | 29.1701 | 3.4308 | -0.031 | -0.0093 | 0.0117 | -0.0344
| SP-19 | 0 | -0.01 | 0.01 | 32.4399 | 29.1075 | 3.1272 | -0.0415 | -0.0123 | 0.0156 | -0.046
| SP-20 | 0 | -0.01 | 0.01 | 32.3537 | 28.9691 | 2.7886 | -0.0226 | -0.0067 | 0.0085 | -0.025
| SP-21 | 0 | -0.01 | 0.01 | 32.8917 | 28.6134 | 3.8873 | -0.028 | -0.0098 | 0.011 | -0.0317
| SP-22 | 0 | -0.01 | 0.01 | 32.8244 | 28.5477 | 3.6563 | -0.0424 | -0.0149 | 0.0166 | -0.0479
| SP-23 | 0 | -0.01 | 0.01 | 32.7313 | 28.4584 | 3.3383 | -0.0476 | -0.0167 | 0.0186 | -0.0538
| SP-24 | 0 | -0.01 | 0.01 | 32.6257 | 28.3285 | 2.9524 | -0.039 | -0.0138 | 0.0153 | -0.0441
Number of points |  |  | 12
Out of Tolerance |  |  | 11
Performance |  |  | 8.333%
| Mean |  | -0.040525
| Std. deviation |  | 0.014446
| Max. Value |  | -0.0088
| Min. Value |  | -0.0607
Inspection Group 48
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 33.3345 | 27.7787 | 4.1926 | -0.0185 | -0.0077 | 0.0076 | -0.0215
| SP-14 | 0 | -0.01 | 0.01 | 33.2585 | 27.6879 | 3.9143 | -0.0372 | -0.0155 | 0.0152 | -0.043
| SP-15 | 0 | -0.01 | 0.01 | 33.1541 | 27.5871 | 3.5563 | -0.0459 | -0.0192 | 0.0188 | -0.0532
| SP-16 | 0 | -0.01 | 0.01 | 33.0638 | 27.5005 | 3.2479 | -0.0518 | -0.0216 | 0.0213 | -0.0601
| SP-17 | 0 | -0.01 | 0.01 | 33.7847 | 27.0308 | 4.4808 | -0.0066 | -0.0031 | 0.0028 | -0.0078
| SP-18 | 0 | -0.01 | 0.01 | 33.6588 | 26.9566 | 4.0998 | -0.0166 | -0.0077 | 0.007 | -0.0196
| SP-19 | 0 | -0.01 | 0.01 | 33.5933 | 26.7999 | 3.7709 | -0.0138 | -0.0065 | 0.0058 | -0.0163
| SP-20 | 0 | -0.01 | 0.01 | 33.5179 | 26.6545 | 3.4307 | -0.0214 | -0.0101 | 0.0091 | -0.0253
| SP-21 | 0 | -0.01 | 0.01 | 34.2613 | 26.2164 | 4.66 | -0.026 | -0.0134 | 0.0112 | -0.0313
| SP-22 | 0 | -0.01 | 0.01 | 34.1589 | 26.0915 | 4.2747 | -0.0346 | -0.0179 | 0.015 | -0.0417
| SP-23 | 0 | -0.01 | 0.01 | 34.076 | 25.9888 | 3.9607 | -0.0264 | -0.0137 | 0.0115 | -0.0319
| SP-24 | 0 | -0.01 | 0.01 | 34.0092 | 25.8612 | 3.6545 | -0.0204 | -0.0106 | 0.0089 | -0.0247
Number of points |  |  | 12
Out of Tolerance |  |  | 11
Performance |  |  | 8.333%
| Mean |  | -0.031367
| Std. deviation |  | 0.015504
| Max. Value |  | -0.0078
| Min. Value |  | -0.0601
Inspection Group 49
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 34.7946 | 25.4474 | 4.9368 | -0.0183 | -0.0102 | 0.0081 | -0.0224
| SP-14 | 0 | -0.01 | 0.01 | 34.713 | 25.3099 | 4.5786 | -0.0258 | -0.0145 | 0.0114 | -0.0317
| SP-15 | 0 | -0.01 | 0.01 | 34.6208 | 25.1957 | 4.2257 | -0.0237 | -0.0133 | 0.0105 | -0.0292
| SP-16 | 0 | -0.01 | 0.01 | 34.5241 | 25.0788 | 3.8594 | -0.0166 | -0.0094 | 0.0074 | -0.0204
| SP-17 | 0 | -0.01 | 0.01 | 35.2734 | 24.7094 | 5.0567 | -0.0022 | -0.0013 | 0.001 | -0.0027
| SP-18 | 0 | -0.01 | 0.01 | 35.1734 | 24.579 | 4.6618 | -0.0056 | -0.0033 | 0.0025 | -0.007
| SP-19 | 0 | -0.01 | 0.01 | 35.125 | 24.4401 | 4.3699 | -0.0083 | -0.005 | 0.0037 | -0.0104
| SP-20 | 0 | -0.01 | 0.01 | 35.0416 | 24.3245 | 4.0313 | 0.0054 | 0.0033 | -0.0024 | 0.0068
| SP-21 | 0 | -0.01 | 0.01 | 35.8506 | 23.8261 | 5.1291 | -0.0148 | -0.0094 | 0.0068 | -0.0188
| SP-22 | 0 | -0.01 | 0.01 | 35.7437 | 23.7588 | 4.8022 | -0.0182 | -0.0116 | 0.0084 | -0.0232
| SP-23 | 0 | -0.01 | 0.01 | 35.6306 | 23.686 | 4.4546 | -0.0181 | -0.0115 | 0.0083 | -0.023
| SP-24 | 0 | -0.01 | 0.01 | 35.5351 | 23.5765 | 4.0943 | -0.0094 | -0.006 | 0.0043 | -0.0119
Number of points |  |  | 12
Out of Tolerance |  |  | 9
Performance |  |  | 25%
| Mean |  | -0.016158
| Std. deviation |  | 0.011343
| Max. Value |  | 0.0068
| Min. Value |  | -0.0317
Inspection Group 50
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 36.4293 | 23.0472 | 5.2757 | -0.0089 | -0.006 | 0.0041 | -0.0115
| SP-14 | 0 | -0.01 | 0.01 | 36.3248 | 22.9586 | 4.9229 | -0.012 | -0.0081 | 0.0056 | -0.0155
| SP-15 | 0 | -0.01 | 0.01 | 36.2224 | 22.8913 | 4.6057 | -0.0142 | -0.0096 | 0.0066 | -0.0183
| SP-16 | 0 | -0.01 | 0.01 | 36.1145 | 22.8207 | 4.272 | -0.0098 | -0.0066 | 0.0046 | -0.0127
| SP-17 | 0 | -0.01 | 0.01 | 36.9912 | 22.271 | 5.3284 | -0.0008 | -0.0006 | 0.0004 | -0.0011
| SP-18 | 0 | -0.01 | 0.01 | 36.9019 | 22.1651 | 4.9799 | 0.0013 | 0.0009 | -0.0006 | 0.0017
| SP-19 | 0 | -0.01 | 0.01 | 36.8339 | 22.0448 | 4.6544 | 0.0102 | 0.0073 | -0.0048 | 0.0135
| SP-20 | 0 | -0.01 | 0.01 | 36.7421 | 21.9391 | 4.3006 | 0.0148 | 0.0106 | -0.007 | 0.0195
| SP-21 | 0 | -0.01 | 0.01 | 37.6414 | 21.4156 | 5.3821 | 0.0006 | 0.0005 | -0.0003 | 0.0008
| SP-22 | 0 | -0.01 | 0.01 | 37.5084 | 21.3508 | 5.0028 | -0.0007 | -0.0006 | 0.0004 | -0.001
| SP-23 | 0 | -0.01 | 0.01 | 37.3908 | 21.2928 | 4.6663 | 0.0077 | 0.0058 | -0.0037 | 0.0103
| SP-24 | 0 | -0.01 | 0.01 | 37.2695 | 21.2133 | 4.2885 | 0.0149 | 0.0112 | -0.0072 | 0.02
Number of points |  |  | 12
Out of Tolerance |  |  | 8
Performance |  |  | 33.33%
| Mean |  | 0.000475
| Std. deviation |  | 0.013281
| Max. Value |  | 0.02
| Min. Value |  | -0.0183
Inspection Group 51
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 38.222 | 20.6267 | 5.3254 | 0.0101 | 0.008 | -0.0049 | 0.0138
| SP-14 | 0 | -0.01 | 0.01 | 38.1043 | 20.5643 | 4.982 | 0.0075 | 0.006 | -0.0037 | 0.0103
| SP-15 | 0 | -0.01 | 0.01 | 38.0042 | 20.5152 | 4.6964 | 0.0043 | 0.0034 | -0.0021 | 0.0059
| SP-16 | 0 | -0.01 | 0.01 | 37.8167 | 20.4757 | 4.2463 | 0.0198 | 0.0156 | -0.0096 | 0.0269
| SP-17 | 0 | -0.01 | 0.01 | 38.8688 | 19.7999 | 5.2755 | 0.0198 | 0.0165 | -0.0098 | 0.0276
| SP-18 | 0 | -0.01 | 0.01 | 38.7258 | 19.7515 | 4.9056 | 0.0162 | 0.0135 | -0.008 | 0.0226
| SP-19 | 0 | -0.01 | 0.01 | 38.627 | 19.671 | 4.571 | 0.0067 | 0.0055 | -0.0033 | 0.0093
| SP-20 | 0 | -0.01 | 0.01 | 38.4857 | 19.6289 | 4.2151 | 0.0203 | 0.0168 | -0.01 | 0.0282
| SP-21 | 0 | -0.01 | 0.01 | 39.5397 | 18.9679 | 5.1923 | 0.0047 | 0.0041 | -0.0024 | 0.0067
| SP-22 | 0 | -0.01 | 0.01 | 39.3489 | 18.9448 | 4.7765 | 0.0019 | 0.0016 | -0.0009 | 0.0027
| SP-23 | 0 | -0.01 | 0.01 | 39.1883 | 18.9501 | 4.4684 | -0.0008 | -0.0007 | 0.0004 | -0.0012
| SP-24 | 0 | -0.01 | 0.01 | 39.0506 | 18.9269 | 4.1561 | 0.0051 | 0.0044 | -0.0026 | 0.0072
Number of points |  |  | 12
Out of Tolerance |  |  | 6
Performance |  |  | 50%
| Mean |  | 0.013333
| Std. deviation |  | 0.010363
| Max. Value |  | 0.0282
| Min. Value |  | -0.0012
Inspection Group 52
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 40.1328 | 18.1609 | 4.9324 | -0.0106 | -0.0098 | 0.0056 | -0.0155
| SP-14 | 0 | -0.01 | 0.01 | 40.0422 | 18.0588 | 4.5789 | -0.0064 | -0.0059 | 0.0034 | -0.0093
| SP-15 | 0 | -0.01 | 0.01 | 39.9346 | 18.0023 | 4.2736 | 0.0003 | 0.0003 | -0.0002 | 0.0004
| SP-16 | 0 | -0.01 | 0.01 | 39.7587 | 17.9821 | 3.9011 | 0.0008 | 0.0007 | -0.0004 | 0.0012
| SP-17 | 0 | -0.01 | 0.01 | 40.796 | 17.3488 | 4.7368 | -0.013 | -0.0126 | 0.007 | -0.0194
| SP-18 | 0 | -0.01 | 0.01 | 40.5682 | 17.3656 | 4.343 | -0.016 | -0.0153 | 0.0086 | -0.0237
| SP-19 | 0 | -0.01 | 0.01 | 40.4349 | 17.3404 | 4.0498 | 0.0018 | 0.0017 | -0.0009 | 0.0026
| SP-20 | 0 | -0.01 | 0.01 | 40.2972 | 17.3093 | 3.738 | 0.0033 | 0.0031 | -0.0017 | 0.0048
| SP-21 | 0 | -0.01 | 0.01 | 41.3422 | 16.5889 | 4.3668 | -0.0175 | -0.0175 | 0.0097 | -0.0265
| SP-22 | 0 | -0.01 | 0.01 | 41.1868 | 16.5774 | 4.0657 | -0.0151 | -0.015 | 0.0084 | -0.0229
| SP-23 | 0 | -0.01 | 0.01 | 41.0382 | 16.5628 | 3.7716 | -0.0086 | -0.0085 | 0.0048 | -0.013
| SP-24 | 0 | -0.01 | 0.01 | 40.8482 | 16.5432 | 3.3939 | 0.0034 | 0.0033 | -0.0019 | 0.0051
Number of points |  |  | 12
Out of Tolerance |  |  | 6
Performance |  |  | 50%
| Mean |  | -0.0096833
| Std. deviation |  | 0.012033
| Max. Value |  | 0.0051
| Min. Value |  | -0.0265
Inspection Group 53
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 41.9315 | 15.8525 | 4.0806 | -0.0181 | -0.0188 | 0.0103 | -0.0281
| SP-14 | 0 | -0.01 | 0.01 | 41.7793 | 15.8363 | 3.7849 | -0.0163 | -0.0168 | 0.0093 | -0.0252
| SP-15 | 0 | -0.01 | 0.01 | 41.6066 | 15.8194 | 3.4523 | -0.015 | -0.0153 | 0.0086 | -0.0231
| SP-16 | 0 | -0.01 | 0.01 | 41.3684 | 15.8322 | 3.0587 | 0.0058 | 0.0058 | -0.0033 | 0.0089
| SP-17 | 0 | -0.01 | 0.01 | 42.5569 | 14.9898 | 3.5927 | -0.0057 | -0.0061 | 0.0033 | -0.009
| SP-18 | 0 | -0.01 | 0.01 | 42.2734 | 15.0753 | 3.2663 | -0.0023 | -0.0024 | 0.0014 | -0.0036
| SP-19 | 0 | -0.01 | 0.01 | 42.0664 | 15.103 | 2.9646 | -0.0061 | -0.0064 | 0.0036 | -0.0095
| SP-20 | 0 | -0.01 | 0.01 | 41.8908 | 15.0815 | 2.6292 | 0.0099 | 0.0103 | -0.0059 | 0.0155
| SP-21 | 0 | -0.01 | 0.01 | 43.1736 | 14.0929 | 2.9981 | -0.0113 | -0.0125 | 0.0069 | -0.0182
| SP-22 | 0 | -0.01 | 0.01 | 42.9115 | 14.1548 | 2.68 | -0.0106 | -0.0116 | 0.0065 | -0.017
| SP-23 | 0 | -0.01 | 0.01 | 42.664 | 14.2051 | 2.365 | -0.0098 | -0.0106 | 0.006 | -0.0156
| SP-24 | 0 | -0.01 | 0.01 | 42.4375 | 14.2349 | 2.0478 | -0.0032 | -0.0035 | 0.002 | -0.0051
Number of points |  |  | 12
Out of Tolerance |  |  | 7
Performance |  |  | 41.67%
| Mean |  | -0.010833
| Std. deviation |  | 0.013291
| Max. Value |  | 0.0155
| Min. Value |  | -0.0281
Inspection Group 54
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-12 | 0 | -0.01 | 0.01 | 43.6183 | 13.4046 | 2.4834 | -0.0194 | -0.0217 | 0.0122 | -0.0316
| SP-13 | 0 | -0.01 | 0.01 | 43.3956 | 13.4452 | 2.2011 | -0.0212 | -0.0236 | 0.0133 | -0.0344
| SP-14 | 0 | -0.01 | 0.01 | 43.1849 | 13.4856 | 1.9373 | -0.0152 | -0.0168 | 0.0096 | -0.0246
| SP-15 | 0 | -0.01 | 0.01 | 42.9132 | 13.5428 | 1.6065 | -0.0095 | -0.0104 | 0.006 | -0.0154
| SP-16 | 0 | -0.01 | 0.01 | 44.0417 | 12.754 | 1.9976 | -0.0076 | -0.0087 | 0.0049 | -0.0126
| SP-17 | 0 | -0.01 | 0.01 | 43.7986 | 12.8023 | 1.7048 | -0.0176 | -0.0198 | 0.0114 | -0.0289
| SP-18 | 0 | -0.01 | 0.01 | 43.5542 | 12.8635 | 1.4326 | -0.0116 | -0.0129 | 0.0075 | -0.0189
| SP-19 | 0 | -0.01 | 0.01 | 43.2464 | 12.9485 | 1.1033 | -0.0055 | -0.0061 | 0.0035 | -0.0089
| SP-20 | 0 | -0.01 | 0.01 | 44.5024 | 11.9695 | 1.3363 | -0.0104 | -0.0118 | 0.0069 | -0.0172
| SP-21 | 0 | -0.01 | 0.01 | 44.2206 | 12.0526 | 1.0537 | -0.0234 | -0.0266 | 0.0155 | -0.0387
| SP-22 | 0 | -0.01 | 0.01 | 43.9727 | 12.1447 | 0.8372 | -0.0245 | -0.0277 | 0.0162 | -0.0403
| SP-23 | 0 | -0.01 | 0.01 | 43.7102 | 12.1978 | 0.532 | -0.0171 | -0.0192 | 0.0113 | -0.0281
Number of points |  |  | 12
Out of Tolerance |  |  | 11
Performance |  |  | 8.333%
| Mean |  | -0.024967
| Std. deviation |  | 0.010365
| Max. Value |  | -0.0089
| Min. Value |  | -0.0403
Inspection Group 55
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 44.9536 | 11.2783 | 0.8276 | -0.0223 | -0.0256 | 0.0151 | -0.0371
| SP-14 | 0 | -0.01 | 0.01 | 44.7065 | 11.3308 | 0.5508 | -0.0307 | -0.0352 | 0.0208 | -0.0511
| SP-15 | 0 | -0.01 | 0.01 | 44.4546 | 11.3894 | 0.2771 | -0.0346 | -0.0395 | 0.0234 | -0.0575
| SP-16 | 0 | -0.01 | 0.01 | 44.149 | 11.4579 | -0.0595 | -0.0288 | -0.0327 | 0.0195 | -0.0477
| SP-17 | 0 | -0.01 | 0.01 | 45.3255 | 10.7131 | 0.4233 | -0.025 | -0.0284 | 0.0171 | -0.0415
| SP-18 | 0 | -0.01 | 0.01 | 45.0633 | 10.7825 | 0.1546 | -0.0285 | -0.0324 | 0.0194 | -0.0473
| SP-19 | 0 | -0.01 | 0.01 | 44.7842 | 10.8528 | -0.1372 | -0.0254 | -0.0289 | 0.0173 | -0.0422
| SP-20 | 0 | -0.01 | 0.01 | 44.5215 | 10.909 | -0.4287 | -0.0266 | -0.0302 | 0.0181 | -0.0441
| SP-21 | 0 | -0.01 | 0.01 | 45.8174 | 9.9823 | -0.0537 | -0.0456 | -0.0499 | 0.0309 | -0.0743
| SP-22 | 0 | -0.01 | 0.01 | 45.5466 | 10.0448 | -0.352 | -0.042 | -0.0464 | 0.0285 | -0.0687
| SP-23 | 0 | -0.01 | 0.01 | 45.3394 | 10.0848 | -0.5924 | -0.0376 | -0.0418 | 0.0255 | -0.0617
| SP-24 | 0 | -0.01 | 0.01 | 45.0887 | 10.13 | -0.8883 | -0.0334 | -0.0373 | 0.0226 | -0.0549
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.052342
| Std. deviation |  | 0.01144
| Max. Value |  | -0.0371
| Min. Value |  | -0.0743
Inspection Group 56
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 46.3237 | 9.2547 | -0.4509 | -0.0522 | -0.0535 | 0.0339 | -0.0821
| SP-14 | 0 | -0.01 | 0.01 | 46.124 | 9.2726 | -0.7303 | -0.0477 | -0.0497 | 0.0309 | -0.0755
| SP-15 | 0 | -0.01 | 0.01 | 45.8873 | 9.3123 | -1.032 | -0.0404 | -0.0428 | 0.0261 | -0.0644
| SP-16 | 0 | -0.01 | 0.01 | 45.6181 | 9.3575 | -1.3746 | -0.0331 | -0.0359 | 0.0213 | -0.0533
| SP-17 | 0 | -0.01 | 0.01 | 46.774 | 8.6628 | -0.6675 | -0.0507 | -0.0487 | 0.031 | -0.0768
| SP-18 | 0 | -0.01 | 0.01 | 46.5603 | 8.6796 | -0.9921 | -0.043 | -0.0424 | 0.0261 | -0.0658
| SP-19 | 0 | -0.01 | 0.01 | 46.378 | 8.6781 | -1.2973 | -0.0348 | -0.0351 | 0.0209 | -0.0537
| SP-20 | 0 | -0.01 | 0.01 | 46.1434 | 8.7001 | -1.6533 | -0.0218 | -0.0226 | 0.0129 | -0.0339
| SP-21 | 0 | -0.01 | 0.01 | 47.4246 | 7.759 | -0.9807 | -0.0785 | -0.0691 | 0.0414 | -0.1125
| SP-22 | 0 | -0.01 | 0.01 | 47.2606 | 7.7793 | -1.2585 | -0.065 | -0.059 | 0.0339 | -0.0942
| SP-23 | 0 | -0.01 | 0.01 | 47.0395 | 7.8039 | -1.6421 | -0.047 | -0.0445 | 0.0241 | -0.069
| SP-24 | 0 | -0.01 | 0.01 | 46.8524 | 7.836 | -1.9488 | -0.0242 | -0.0236 | 0.0122 | -0.0359
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.068092
| Std. deviation |  | 0.022627
| Max. Value |  | -0.0339
| Min. Value |  | -0.1125
Inspection Group 57
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 48.0198 | 6.9731 | -1.1228 | -0.0779 | -0.0634 | 0.0352 | -0.1064
| SP-14 | 0 | -0.01 | 0.01 | 47.8454 | 6.9967 | -1.4697 | -0.0683 | -0.058 | 0.0301 | -0.0945
| SP-15 | 0 | -0.01 | 0.01 | 47.7038 | 7.0222 | -1.7433 | -0.0631 | -0.0554 | 0.0274 | -0.0883
| SP-16 | 0 | -0.01 | 0.01 | 47.4905 | 7.0684 | -2.1442 | -0.0331 | -0.0305 | 0.014 | -0.0472
| SP-17 | 0 | -0.01 | 0.01 | 48.5618 | 6.2059 | -1.2691 | -0.0735 | -0.0558 | 0.0276 | -0.0963
| SP-18 | 0 | -0.01 | 0.01 | 48.4191 | 6.2175 | -1.6325 | -0.0647 | -0.0514 | 0.0233 | -0.0859
| SP-19 | 0 | -0.01 | 0.01 | 48.2266 | 6.3184 | -1.9409 | -0.0417 | -0.0347 | 0.0149 | -0.0563
| SP-20 | 0 | -0.01 | 0.01 | 48.0997 | 6.322 | -2.295 | -0.0116 | -0.01 | 0.004 | -0.0158
| SP-21 | 0 | -0.01 | 0.01 | 49.1678 | 5.299 | -1.4366 | -0.0855 | -0.0599 | 0.0245 | -0.1072
| SP-22 | 0 | -0.01 | 0.01 | 49.0758 | 5.3004 | -1.7631 | -0.075 | -0.0547 | 0.0203 | -0.0951
| SP-23 | 0 | -0.01 | 0.01 | 48.9749 | 5.3025 | -2.1427 | -0.0557 | -0.0425 | 0.0141 | -0.0715
| SP-24 | 0 | -0.01 | 0.01 | 48.8564 | 5.3487 | -2.4775 | -0.0331 | -0.0264 | 0.0079 | -0.0431
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.075633
| Std. deviation |  | 0.028934
| Max. Value |  | -0.0158
| Min. Value |  | -0.1072
Inspection Group 58
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 49.694 | 4.4871 | -1.5144 | -0.0656 | -0.0421 | 0.0142 | -0.0792
| SP-14 | 0 | -0.01 | 0.01 | 49.62 | 4.4853 | -1.8747 | -0.0648 | -0.0433 | 0.013 | -0.079
| SP-15 | 0 | -0.01 | 0.01 | 49.5231 | 4.5037 | -2.3196 | -0.0437 | -0.0307 | 0.0078 | -0.0539
| SP-16 | 0 | -0.01 | 0.01 | 49.4485 | 4.5264 | -2.6585 | -0.0274 | -0.0199 | 0.0045 | -0.0342
| SP-17 | 0 | -0.01 | 0.01 | 50.3778 | 3.2828 | -1.7191 | -0.0434 | -0.0238 | 0.0057 | -0.0498
| SP-18 | 0 | -0.01 | 0.01 | 50.3153 | 3.3151 | -2.0739 | -0.0397 | -0.0227 | 0.0047 | -0.046
| SP-19 | 0 | -0.01 | 0.01 | 50.235 | 3.3794 | -2.4572 | -0.0274 | -0.0164 | 0.0029 | -0.0321
| SP-20 | 0 | -0.01 | 0.01 | 50.1602 | 3.4396 | -2.8475 | 0.01 | 0.0062 | -0.0009 | 0.0118
| SP-21 | 0 | -0.01 | 0.01 | 50.8765 | 2.2531 | -1.9297 | -0.0545 | -0.025 | 0.0044 | -0.0602
| SP-22 | 0 | -0.01 | 0.01 | 50.8153 | 2.3266 | -2.2788 | -0.0445 | -0.0212 | 0.0032 | -0.0494
| SP-23 | 0 | -0.01 | 0.01 | 50.7547 | 2.3994 | -2.6553 | -0.0359 | -0.0177 | 0.0022 | -0.0401
| SP-24 | 0 | -0.01 | 0.01 | 50.7006 | 2.4652 | -3.0186 | -0.0063 | -0.0032 | 0.0003 | -0.0071
Number of points |  |  | 12
Out of Tolerance |  |  | 11
Performance |  |  | 8.333%
| Mean |  | -0.043267
| Std. deviation |  | 0.026308
| Max. Value |  | 0.0118
| Min. Value |  | -0.0792
Inspection Group 59
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 51.2226 | 1.3862 | -2.1293 | -0.042 | -0.0156 | 0.0025 | -0.0449
| SP-14 | 0 | -0.01 | 0.01 | 51.1836 | 1.4403 | -2.4679 | -0.0369 | -0.0141 | 0.0019 | -0.0395
| SP-15 | 0 | -0.01 | 0.01 | 51.1336 | 1.5222 | -2.8493 | -0.0203 | -0.008 | 0.0009 | -0.0219
| SP-16 | 0 | -0.01 | 0.01 | 51.0743 | 1.6264 | -3.2849 | -0.0102 | -0.0042 | 0.0004 | -0.0111
| SP-17 | 0 | -0.01 | 0.01 | 51.6159 | 0.0328 | -2.3671 | -0.0408 | -0.0092 | 0.002 | -0.0418
| SP-18 | 0 | -0.01 | 0.01 | 51.5781 | 0.1249 | -2.7185 | -0.0412 | -0.0096 | 0.0019 | -0.0424
| SP-19 | 0 | -0.01 | 0.01 | 51.5404 | 0.2141 | -3.1015 | -0.0174 | -0.0042 | 0.0007 | -0.018
| SP-20 | 0 | -0.01 | 0.01 | 51.5021 | 0.2994 | -3.5485 | -0.0181 | -0.0045 | 0.0007 | -0.0187
| SP-21 | 0 | -0.01 | 0.01 | 51.7959 | -1.4296 | -2.768 | -0.0802 | -0.0035 | 0.0052 | -0.0805
| SP-22 | 0 | -0.01 | 0.01 | 51.774 | -1.396 | -3.0836 | -0.0827 | -0.0035 | 0.0053 | -0.083
| SP-23 | 0 | -0.01 | 0.01 | 51.7469 | -1.3319 | -3.4662 | -0.0911 | -0.0039 | 0.0058 | -0.0914
| SP-24 | 0 | -0.01 | 0.01 | 51.7121 | -1.1479 | -3.8831 | -0.0489 | -0.0028 | 0.0029 | -0.0491
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.045192
| Std. deviation |  | 0.027009
| Max. Value |  | -0.0111
| Min. Value |  | -0.0914
Inspection Group 60
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 51.7462 | -2.5072 | -3.0258 | -0.0704 | 0.0069 | 0.0062 | -0.0711
| SP-14 | 0 | -0.01 | 0.01 | 51.7205 | -2.4731 | -3.3577 | -0.0743 | 0.0076 | 0.0065 | -0.075
| SP-15 | 0 | -0.01 | 0.01 | 51.6943 | -2.3933 | -3.7479 | -0.0735 | 0.0074 | 0.0064 | -0.0741
| SP-16 | 0 | -0.01 | 0.01 | 51.6691 | -2.3094 | -4.1337 | -0.0538 | 0.0054 | 0.0047 | -0.0543
| SP-17 | 0 | -0.01 | 0.01 | 51.4332 | -4.0129 | -3.2893 | -0.0673 | 0.0187 | 0.0082 | -0.0704
| SP-18 | 0 | -0.01 | 0.01 | 51.3979 | -3.9679 | -3.6806 | -0.0541 | 0.0152 | 0.0066 | -0.0566
| SP-19 | 0 | -0.01 | 0.01 | 51.3736 | -3.8694 | -4.1042 | -0.0221 | 0.0062 | 0.0027 | -0.0231
| SP-20 | 0 | -0.01 | 0.01 | 51.348 | -3.8197 | -4.428 | -0.0094 | 0.0026 | 0.0011 | -0.0098
| SP-21 | 0 | -0.01 | 0.01 | 50.9932 | -5.2045 | -3.572 | -0.0697 | 0.0275 | 0.0099 | -0.0756
| SP-22 | 0 | -0.01 | 0.01 | 50.9639 | -5.1574 | -3.9099 | -0.0589 | 0.0232 | 0.0083 | -0.0638
| SP-23 | 0 | -0.01 | 0.01 | 50.9721 | -5.0077 | -4.2678 | -0.022 | 0.0085 | 0.0031 | -0.0238
| SP-24 | 0 | -0.01 | 0.01 | 50.9573 | -4.9033 | -4.6653 | -0.0137 | 0.0052 | 0.0019 | -0.0148
Number of points |  |  | 12
Out of Tolerance |  |  | 11
Performance |  |  | 8.333%
| Mean |  | -0.051033
| Std. deviation |  | 0.025633
| Max. Value |  | -0.0098
| Min. Value |  | -0.0756
Inspection Group 61
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 50.4708 | -6.2816 | -3.9012 | -0.0573 | 0.0277 | 0.0086 | -0.0642
| SP-14 | 0 | -0.01 | 0.01 | 50.4569 | -6.1941 | -4.2766 | -0.0482 | 0.023 | 0.0071 | -0.0538
| SP-15 | 0 | -0.01 | 0.01 | 50.4457 | -6.109 | -4.6333 | -0.0403 | 0.019 | 0.0058 | -0.045
| SP-16 | 0 | -0.01 | 0.01 | 50.468 | -5.9552 | -4.9852 | -0.0205 | 0.0095 | 0.0028 | -0.0228
| SP-17 | 0 | -0.01 | 0.01 | 49.9367 | -7.2413 | -4.1861 | -0.042 | 0.0226 | 0.0061 | -0.0481
| SP-18 | 0 | -0.01 | 0.01 | 49.9583 | -7.0986 | -4.5687 | -0.0406 | 0.0215 | 0.0057 | -0.0462
| SP-19 | 0 | -0.01 | 0.01 | 49.9902 | -6.941 | -4.9382 | -0.0424 | 0.022 | 0.0057 | -0.0481
| SP-20 | 0 | -0.01 | 0.01 | 50.0388 | -6.7576 | -5.2842 | -0.0201 | 0.0102 | 0.0026 | -0.0226
| SP-21 | 0 | -0.01 | 0.01 | 49.2733 | -8.2899 | -4.7793 | -0.0758 | 0.0432 | 0.0091 | -0.0877
| SP-22 | 0 | -0.01 | 0.01 | 49.3308 | -8.1191 | -5.112 | -0.0765 | 0.0429 | 0.0088 | -0.0881
| SP-23 | 0 | -0.01 | 0.01 | 49.3884 | -7.9475 | -5.4503 | -0.0754 | 0.0414 | 0.0081 | -0.0864
| SP-24 | 0 | -0.01 | 0.01 | 49.4405 | -7.785 | -5.8003 | -0.0563 | 0.0303 | 0.0056 | -0.0642
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.056433
| Std. deviation |  | 0.022676
| Max. Value |  | -0.0226
| Min. Value |  | -0.0881
Inspection Group 62
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 48.7989 | -9.0244 | -5.2112 | -0.0705 | 0.0418 | 0.007 | -0.0823
| SP-14 | 0 | -0.01 | 0.01 | 48.8653 | -8.8513 | -5.5821 | -0.072 | 0.0418 | 0.0065 | -0.0835
| SP-15 | 0 | -0.01 | 0.01 | 48.9154 | -8.7258 | -5.839 | -0.0674 | 0.0385 | 0.0056 | -0.0779
| SP-16 | 0 | -0.01 | 0.01 | 49.0011 | -8.5195 | -6.2419 | -0.0686 | 0.0382 | 0.0048 | -0.0786
| SP-17 | 0 | -0.01 | 0.01 | 48.3316 | -9.7332 | -5.6171 | -0.0598 | 0.0373 | 0.0047 | -0.0706
| SP-18 | 0 | -0.01 | 0.01 | 48.415 | -9.5522 | -6.0129 | -0.0668 | 0.0408 | 0.0043 | -0.0784
| SP-19 | 0 | -0.01 | 0.01 | 48.512 | -9.3611 | -6.3184 | -0.041 | 0.0244 | 0.0022 | -0.0478
| SP-20 | 0 | -0.01 | 0.01 | 48.5929 | -9.1977 | -6.6372 | -0.0481 | 0.028 | 0.002 | -0.0557
| SP-21 | 0 | -0.01 | 0.01 | 47.765 | -10.5587 | -6.1747 | -0.0667 | 0.0444 | 0.0029 | -0.0801
| SP-22 | 0 | -0.01 | 0.01 | 47.953 | -10.2578 | -6.3955 | -0.0642 | 0.0414 | 0.0024 | -0.0764
| SP-23 | 0 | -0.01 | 0.01 | 48.084 | -10.0327 | -6.7922 | -0.0492 | 0.0308 | 0.0011 | -0.058
| SP-24 | 0 | -0.01 | 0.01 | 48.1791 | -9.8721 | -7.0515 | -0.0391 | 0.024 | 0.0005 | -0.0459
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.0696
| Std. deviation |  | 0.01383
| Max. Value |  | -0.0459
| Min. Value |  | -0.0835
Inspection Group 63
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 47.2455 | -11.2977 | -6.7551 | -0.0394 | 0.0277 | 0.0001 | -0.0482
| SP-14 | 0 | -0.01 | 0.01 | 47.3643 | -11.1278 | -7.017 | -0.0412 | 0.0283 | -0.0003 | -0.05
| SP-15 | 0 | -0.01 | 0.01 | 47.4936 | -10.9434 | -7.323 | -0.0481 | 0.0322 | -0.0009 | -0.0579
| SP-16 | 0 | -0.01 | 0.01 | 47.6053 | -10.7885 | -7.6656 | -0.0335 | 0.0219 | -0.0011 | -0.0401
| SP-17 | 0 | -0.01 | 0.01 | 46.842 | -11.8622 | -7.0272 | -0.0092 | 0.0068 | -0.0002 | -0.0115
| SP-18 | 0 | -0.01 | 0.01 | 46.9695 | -11.7035 | -7.4084 | -0.0232 | 0.0166 | -0.0009 | -0.0286
| SP-19 | 0 | -0.01 | 0.01 | 47.0359 | -11.6284 | -7.6892 | -0.0287 | 0.0202 | -0.0015 | -0.0351
| SP-20 | 0 | -0.01 | 0.01 | 47.2181 | -11.3937 | -8.0272 | -0.0104 | 0.0071 | -0.0007 | -0.0127
| SP-21 | 0 | -0.01 | 0.01 | 46.365 | -12.5267 | -7.4866 | -0.0131 | 0.0099 | -0.0009 | -0.0164
| SP-22 | 0 | -0.01 | 0.01 | 46.4602 | -12.4286 | -7.7685 | -0.014 | 0.0105 | -0.0011 | -0.0175
| SP-23 | 0 | -0.01 | 0.01 | 46.56 | -12.3279 | -8.0635 | -0.021 | 0.0154 | -0.0019 | -0.0261
| SP-24 | 0 | -0.01 | 0.01 | 46.7264 | -12.1588 | -8.5182 | -0.0155 | 0.011 | -0.0016 | -0.019
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.030258
| Std. deviation |  | 0.015834
| Max. Value |  | -0.0115
| Min. Value |  | -0.0579
Inspection Group 64
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 45.8205 | -13.2727 | -7.8299 | -0.0016 | 0.0012 | -0.0002 | -0.002
| SP-14 | 0 | -0.01 | 0.01 | 45.9339 | -13.1793 | -8.1677 | -0.0103 | 0.0079 | -0.0013 | -0.0131
| SP-15 | 0 | -0.01 | 0.01 | 46.036 | -13.0946 | -8.4599 | -0.0154 | 0.0116 | -0.0021 | -0.0194
| SP-16 | 0 | -0.01 | 0.01 | 46.1322 | -13.0259 | -8.7884 | -0.0343 | 0.0252 | -0.0049 | -0.0429
| SP-17 | 0 | -0.01 | 0.01 | 45.3839 | -13.8828 | -8.1202 | 0.0061 | -0.0047 | 0.0009 | 0.0078
| SP-18 | 0 | -0.01 | 0.01 | 45.4637 | -13.854 | -8.4833 | -0.0079 | 0.0061 | -0.0013 | -0.0101
| SP-19 | 0 | -0.01 | 0.01 | 45.5913 | -13.7529 | -8.793 | 0.0009 | -0.0006 | 0.0001 | 0.0011
| SP-20 | 0 | -0.01 | 0.01 | 45.7637 | -13.5776 | -9.0424 | -0.0223 | 0.0165 | -0.0037 | -0.028
| SP-21 | 0 | -0.01 | 0.01 | 44.7969 | -14.7187 | -8.4548 | -0.009 | 0.0069 | -0.0017 | -0.0114
| SP-22 | 0 | -0.01 | 0.01 | 44.9232 | -14.6345 | -8.7791 | -0.0165 | 0.0124 | -0.0032 | -0.0209
| SP-23 | 0 | -0.01 | 0.01 | 45.0494 | -14.5506 | -9.1047 | -0.022 | 0.0164 | -0.0043 | -0.0278
| SP-24 | 0 | -0.01 | 0.01 | 45.1621 | -14.4819 | -9.4194 | -0.0214 | 0.0158 | -0.0042 | -0.027
Number of points |  |  | 12
Out of Tolerance |  |  | 9
Performance |  |  | 25%
| Mean |  | -0.016142
| Std. deviation |  | 0.014405
| Max. Value |  | 0.0078
| Min. Value |  | -0.0429
Inspection Group 65
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 44.2859 | -15.4898 | -8.7719 | -0.0016 | 0.0012 | -0.0003 | -0.002
| SP-14 | 0 | -0.01 | 0.01 | 44.4306 | -15.4001 | -9.1455 | -0.0159 | 0.0115 | -0.0034 | -0.0199
| SP-15 | 0 | -0.01 | 0.01 | 44.5466 | -15.3274 | -9.4426 | -0.0269 | 0.0195 | -0.0057 | -0.0337
| SP-16 | 0 | -0.01 | 0.01 | 44.6593 | -15.2508 | -9.711 | -0.0207 | 0.015 | -0.0044 | -0.0259
| SP-17 | 0 | -0.01 | 0.01 | 43.8778 | -16.1404 | -9.0182 | 0.0136 | -0.0093 | 0.003 | 0.0168
| SP-18 | 0 | -0.01 | 0.01 | 44.0142 | -16.0301 | -9.2906 | 0.0059 | -0.0041 | 0.0013 | 0.0073
| SP-19 | 0 | -0.01 | 0.01 | 44.1472 | -15.9492 | -9.639 | -0.0032 | 0.0022 | -0.0007 | -0.0039
| SP-20 | 0 | -0.01 | 0.01 | 44.2991 | -15.8174 | -9.9112 | 0.011 | -0.0077 | 0.0024 | 0.0137
| SP-21 | 0 | -0.01 | 0.01 | 43.3681 | -17.0543 | -9.4017 | 0.0157 | -0.0096 | 0.0035 | 0.0187
| SP-22 | 0 | -0.01 | 0.01 | 43.4851 | -16.9784 | -9.7176 | 0.0069 | -0.0043 | 0.0015 | 0.0083
| SP-23 | 0 | -0.01 | 0.01 | 43.6251 | -16.8882 | -10.0973 | 0.0079 | -0.005 | 0.0017 | 0.0095
| SP-24 | 0 | -0.01 | 0.01 | 43.7714 | -16.7465 | -10.3536 | 0.016 | -0.0103 | 0.0035 | 0.0193
Number of points |  |  | 12
Out of Tolerance |  |  | 7
Performance |  |  | 41.67%
| Mean |  | 0.00068333
| Std. deviation |  | 0.018143
| Max. Value |  | 0.0193
| Min. Value |  | -0.0337
Inspection Group 66
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 42.9814 | -17.8719 | -9.7972 | 0.008 | -0.0044 | 0.0017 | 0.0092
| SP-14 | 0 | -0.01 | 0.01 | 43.0838 | -17.8018 | -10.0936 | -0.0064 | 0.0035 | -0.0014 | -0.0075
| SP-15 | 0 | -0.01 | 0.01 | 43.2197 | -17.7071 | -10.4852 | 0.0076 | -0.0043 | 0.0016 | 0.0088
| SP-16 | 0 | -0.01 | 0.01 | 43.2797 | -17.7064 | -10.7707 | 0.0081 | -0.0046 | 0.0017 | 0.0094
| SP-17 | 0 | -0.01 | 0.01 | 42.6185 | -18.8019 | -10.3207 | 0.0169 | -0.008 | 0.0036 | 0.019
| SP-18 | 0 | -0.01 | 0.01 | 42.7784 | -18.6273 | -10.6802 | 0.0058 | -0.0028 | 0.0012 | 0.0065
| SP-19 | 0 | -0.01 | 0.01 | 42.9032 | -18.5025 | -10.9875 | 0.0311 | -0.0153 | 0.0064 | 0.0352
| SP-20 | 0 | -0.01 | 0.01 | 43.0937 | -18.2312 | -11.252 | 0.0332 | -0.0171 | 0.0066 | 0.0379
| SP-21 | 0 | -0.01 | 0.01 | 42.3943 | -19.7278 | -11.1556 | -0.0065 | 0.0026 | -0.0015 | -0.0071
| SP-22 | 0 | -0.01 | 0.01 | 42.546 | -19.5013 | -11.4259 | -0.0098 | 0.004 | -0.0022 | -0.0108
| SP-23 | 0 | -0.01 | 0.01 | 42.6848 | -19.2958 | -11.6739 | 0.0009 | -0.0004 | 0.0002 | 0.001
| SP-24 | 0 | -0.01 | 0.01 | 42.8394 | -19.0589 | -11.931 | 0.0253 | -0.0108 | 0.0053 | 0.028
Number of points |  |  | 12
Out of Tolerance |  |  | 5
Performance |  |  | 58.33%
| Mean |  | 0.0108
| Std. deviation |  | 0.016342
| Max. Value |  | 0.0379
| Min. Value |  | -0.0108
Inspection Group 67
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 42.3365 | -20.3777 | -11.9024 | -0.0305 | 0.01 | -0.0076 | -0.033
| SP-14 | 0 | -0.01 | 0.01 | 42.4669 | -20.172 | -12.1558 | -0.0275 | 0.0092 | -0.0067 | -0.0297
| SP-15 | 0 | -0.01 | 0.01 | 42.6173 | -19.9321 | -12.4469 | -0.0218 | 0.0075 | -0.0051 | -0.0236
| SP-16 | 0 | -0.01 | 0.01 | 42.7171 | -19.7627 | -12.6256 | -0.0075 | 0.0026 | -0.0017 | -0.0082
| SP-17 | 0 | -0.01 | 0.01 | 42.3579 | -21.113 | -12.8053 | 0.0055 | -0.0014 | 0.0015 | 0.0059
| SP-18 | 0 | -0.01 | 0.01 | 42.5017 | -20.7847 | -13.0266 | -0.0015 | 0.0004 | -0.0004 | -0.0016
| SP-19 | 0 | -0.01 | 0.01 | 42.618 | -20.5068 | -13.1885 | 0.0072 | -0.002 | 0.0018 | 0.0077
| SP-20 | 0 | -0.01 | 0.01 | 42.7625 | -20.234 | -13.4709 | 0.022 | -0.0063 | 0.0051 | 0.0235
| SP-21 | 0 | -0.01 | 0.01 | 42.5605 | -21.6442 | -13.9919 | 0.0233 | -0.0046 | 0.0064 | 0.0246
| SP-22 | 0 | -0.01 | 0.01 | 42.6407 | -21.3854 | -14.0972 | 0.0094 | -0.0019 | 0.0025 | 0.0099
| SP-23 | 0 | -0.01 | 0.01 | 42.7516 | -21.0284 | -14.2381 | 0.0098 | -0.0021 | 0.0025 | 0.0104
| SP-24 | 0 | -0.01 | 0.01 | 42.8616 | -20.7103 | -14.4035 | 0.016 | -0.0036 | 0.0038 | 0.0168
Number of points |  |  | 12
Out of Tolerance |  |  | 7
Performance |  |  | 41.67%
| Mean |  | 0.000225
| Std. deviation |  | 0.019857
| Max. Value |  | 0.0246
| Min. Value |  | -0.033
Inspection Group 68
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 42.7264 | -22.0713 | -14.8595 | 0.0275 | -0.0043 | 0.0077 | 0.0289
| SP-14 | 0 | -0.01 | 0.01 | 42.8102 | -21.7686 | -14.9919 | 0.0156 | -0.0025 | 0.0042 | 0.0164
| SP-15 | 0 | -0.01 | 0.01 | 42.8972 | -21.4554 | -15.1294 | 0.0216 | -0.0036 | 0.0055 | 0.0225
| SP-16 | 0 | -0.01 | 0.01 | 43.0016 | -21.0457 | -15.2683 | 0.0225 | -0.004 | 0.0054 | 0.0235
| SP-17 | 0 | -0.01 | 0.01 | 42.9327 | -22.3663 | -15.756 | 0.0487 | -0.0059 | 0.0132 | 0.0508
| SP-18 | 0 | -0.01 | 0.01 | 43.0136 | -21.9874 | -15.8853 | 0.0346 | -0.0044 | 0.0089 | 0.036
| SP-19 | 0 | -0.01 | 0.01 | 43.0741 | -21.6536 | -15.9542 | 0.0372 | -0.005 | 0.0091 | 0.0386
| SP-20 | 0 | -0.01 | 0.01 | 43.1563 | -21.248 | -16.0666 | 0.0443 | -0.0063 | 0.0101 | 0.0459
| SP-21 | 0 | -0.01 | 0.01 | 43.2409 | -22.5678 | -17.058 | 0.019 | -0.0016 | 0.0044 | 0.0195
| SP-22 | 0 | -0.01 | 0.01 | 43.2832 | -22.2458 | -17.1237 | 0.0176 | -0.0015 | 0.0039 | 0.0181
| SP-23 | 0 | -0.01 | 0.01 | 43.338 | -21.8394 | -17.2081 | 0.0285 | -0.0027 | 0.0059 | 0.0293
| SP-24 | 0 | -0.01 | 0.01 | 43.387 | -21.4993 | -17.2894 | 0.0406 | -0.004 | 0.008 | 0.0416
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | 0.030925
| Std. deviation |  | 0.011508
| Max. Value |  | 0.0508
| Min. Value |  | 0.0164
Inspection Group 69
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 43.4381 | -22.7375 | -18.004 | -0.0234 | 0.0011 | -0.0048 | -0.0239
| SP-14 | 0 | -0.01 | 0.01 | 43.4691 | -22.4096 | -18.0751 | -0.0046 | 0.0003 | -0.0009 | -0.0047
| SP-15 | 0 | -0.01 | 0.01 | 43.5042 | -22.0506 | -18.1518 | 0.0104 | -0.0006 | 0.0018 | 0.0106
| SP-16 | 0 | -0.01 | 0.01 | 43.5391 | -21.6357 | -18.1955 | 0.032 | -0.0022 | 0.0053 | 0.0325
| SP-17 | 0 | -0.01 | 0.01 | 43.5856 | -22.9126 | -18.8148 | -0.0145 | 0.0002 | -0.0024 | -0.0147
| SP-18 | 0 | -0.01 | 0.01 | 43.6194 | -22.581 | -18.991 | -0.006 | 0.0001 | -0.0009 | -0.0061
| SP-19 | 0 | -0.01 | 0.01 | 43.6329 | -22.2458 | -19.0195 | 0.0214 | -0.0007 | 0.0029 | 0.0216
| SP-20 | 0 | -0.01 | 0.01 | 43.6794 | -21.8092 | -19.264 | 0.0258 | -0.001 | 0.0029 | 0.026
| SP-21 | 0 | -0.01 | 0.01 | 43.7059 | -23.1422 | -19.795 | -0.0142 | 0 | -0.0011 | -0.0142
| SP-22 | 0 | -0.01 | 0.01 | 43.7202 | -22.9024 | -19.9771 | -0.0242 | 0.0002 | -0.0016 | -0.0243
| SP-23 | 0 | -0.01 | 0.01 | 43.7364 | -22.4685 | -20.1488 | -0.0004 | 0 | 0 | -0.0004
| SP-24 | 0 | -0.01 | 0.01 | 43.7527 | -22.0245 | -20.274 | 0.007 | -0.0002 | 0.0003 | 0.007
Number of points |  |  | 12
Out of Tolerance |  |  | 8
Performance |  |  | 33.33%
| Mean |  | 0.00078333
| Std. deviation |  | 0.019049
| Max. Value |  | 0.0325
| Min. Value |  | -0.0243
Inspection Group 70
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 43.7354 | -23.6334 | -20.7614 | -0.0039 | 0.0001 | 0.0001 | -0.0039
| SP-14 | 0 | -0.01 | 0.01 | 43.7363 | -23.3942 | -20.9787 | -0.0077 | 0.0002 | 0.0002 | -0.0077
| SP-15 | 0 | -0.01 | 0.01 | 43.7426 | -22.983 | -21.1428 | -0.0019 | 0 | 0 | -0.0019
| SP-16 | 0 | -0.01 | 0.01 | 43.7454 | -22.6883 | -21.3718 | 0.0015 | 0 | 0 | 0.0015
| SP-17 | 0 | -0.01 | 0.01 | 43.7106 | -23.9831 | -21.201 | 0.0312 | -0.0012 | -0.0017 | 0.0313
| SP-18 | 0 | -0.01 | 0.01 | 43.7079 | -23.7101 | -21.4487 | 0.0164 | -0.0007 | -0.0009 | 0.0164
| SP-19 | 0 | -0.01 | 0.01 | 43.7008 | -23.4975 | -21.7273 | 0.0233 | -0.0011 | -0.0014 | 0.0234
| SP-20 | 0 | -0.01 | 0.01 | 43.6998 | -23.1768 | -21.9851 | 0.0368 | -0.0017 | -0.0023 | 0.0369
| SP-21 | 0 | -0.01 | 0.01 | 43.6753 | -24.3078 | -21.4911 | 0.0568 | -0.0033 | -0.0043 | 0.057
| SP-22 | 0 | -0.01 | 0.01 | 43.6598 | -24.1198 | -21.8375 | 0.0229 | -0.0014 | -0.0018 | 0.0231
| SP-23 | 0 | -0.01 | 0.01 | 43.6473 | -23.9233 | -22.1519 | 0.0282 | -0.0018 | -0.0022 | 0.0283
| SP-24 | 0 | -0.01 | 0.01 | 43.6396 | -23.6997 | -22.4262 | 0.042 | -0.0027 | -0.0033 | 0.0422
Number of points |  |  | 12
Out of Tolerance |  |  | 8
Performance |  |  | 33.33%
| Mean |  | 0.02055
| Std. deviation |  | 0.020304
| Max. Value |  | 0.057
| Min. Value |  | -0.0077
Inspection Group 71
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 43.5712 | -25.1648 | -21.9209 | 0.0606 | -0.0059 | -0.0058 | 0.0611
| SP-14 | 0 | -0.01 | 0.01 | 43.5383 | -25.0651 | -22.3672 | 0.0221 | -0.0021 | -0.0021 | 0.0223
| SP-15 | 0 | -0.01 | 0.01 | 43.5168 | -24.9565 | -22.7031 | 0.0276 | -0.0026 | -0.0026 | 0.0278
| SP-16 | 0 | -0.01 | 0.01 | 43.495 | -24.8427 | -23.0485 | 0.0395 | -0.0036 | -0.0037 | 0.0398
| SP-17 | 0 | -0.01 | 0.01 | 43.4996 | -25.741 | -22.019 | 0.0667 | -0.0076 | -0.0062 | 0.0674
| SP-18 | 0 | -0.01 | 0.01 | 43.4632 | -25.6876 | -22.4774 | 0.0432 | -0.0047 | -0.004 | 0.0437
| SP-19 | 0 | -0.01 | 0.01 | 43.4357 | -25.6675 | -22.8015 | 0.0489 | -0.0052 | -0.0045 | 0.0494
| SP-20 | 0 | -0.01 | 0.01 | 43.3946 | -25.6794 | -23.2384 | 0.0561 | -0.0058 | -0.0051 | 0.0566
| SP-21 | 0 | -0.01 | 0.01 | 43.4426 | -26.158 | -22.1303 | 0.0161 | -0.0018 | -0.0014 | 0.0163
| SP-22 | 0 | -0.01 | 0.01 | 43.4001 | -26.256 | -22.486 | 0.0114 | -0.0013 | -0.001 | 0.0115
| SP-23 | 0 | -0.01 | 0.01 | 43.3532 | -26.3723 | -22.8703 | 0.0147 | -0.0017 | -0.0013 | 0.0149
| SP-24 | 0 | -0.01 | 0.01 | 43.3112 | -26.4828 | -23.2077 | 0.0197 | -0.0022 | -0.0017 | 0.0199
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | 0.035892
| Std. deviation |  | 0.019636
| Max. Value |  | 0.0674
| Min. Value |  | 0.0115
Inspection Group 72
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 43.3343 | -27.1472 | -21.8666 | 0.0123 | -0.0019 | -0.0009 | 0.0125
| SP-14 | 0 | -0.01 | 0.01 | 43.3004 | -27.1978 | -22.2283 | -0.0151 | 0.0023 | 0.0011 | -0.0154
| SP-15 | 0 | -0.01 | 0.01 | 43.26 | -27.3019 | -22.5666 | -0.0157 | 0.0023 | 0.0012 | -0.016
| SP-16 | 0 | -0.01 | 0.01 | 43.2077 | -27.4227 | -23.0318 | 0.012 | -0.0017 | -0.0009 | 0.0121
| SP-17 | 0 | -0.01 | 0.01 | 43.2467 | -27.7713 | -21.5026 | -0.0027 | 0.0005 | 0.0001 | -0.0028
| SP-18 | 0 | -0.01 | 0.01 | 43.1897 | -27.9911 | -21.8561 | -0.0177 | 0.0033 | 0.0008 | -0.018
| SP-19 | 0 | -0.01 | 0.01 | 43.1444 | -28.1587 | -22.1748 | 0.0011 | -0.0002 | -0.0001 | 0.0011
| SP-20 | 0 | -0.01 | 0.01 | 43.0833 | -28.4002 | -22.5563 | 0.0292 | -0.0051 | -0.0014 | 0.0296
| SP-21 | 0 | -0.01 | 0.01 | 43.1562 | -28.269 | -21.1616 | -0.0236 | 0.005 | 0.0004 | -0.0241
| SP-22 | 0 | -0.01 | 0.01 | 43.0849 | -28.5807 | -21.4586 | -0.0369 | 0.0078 | 0.0006 | -0.0377
| SP-23 | 0 | -0.01 | 0.01 | 43.0289 | -28.8271 | -21.7138 | -0.0396 | 0.0083 | 0.0007 | -0.0404
| SP-24 | 0 | -0.01 | 0.01 | 42.9574 | -29.1353 | -22.1449 | -0.0047 | 0.001 | 0.0001 | -0.0048
Number of points |  |  | 12
Out of Tolerance |  |  | 9
Performance |  |  | 25%
| Mean |  | -0.0086583
| Std. deviation |  | 0.020808
| Max. Value |  | 0.0296
| Min. Value |  | -0.0404
Inspection Group 73
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 42.9611 | -29.0899 | -20.5096 | -0.0385 | 0.0097 | -0.0012 | -0.0397
| SP-14 | 0 | -0.01 | 0.01 | 42.916 | -29.306 | -20.8151 | -0.0507 | 0.0125 | -0.0014 | -0.0523
| SP-15 | 0 | -0.01 | 0.01 | 42.8607 | -29.5582 | -21.0721 | -0.0468 | 0.0115 | -0.0012 | -0.0482
| SP-16 | 0 | -0.01 | 0.01 | 42.7908 | -29.8794 | -21.4167 | -0.0282 | 0.0068 | -0.0006 | -0.029
| SP-17 | 0 | -0.01 | 0.01 | 42.7377 | -29.728 | -19.7163 | -0.0392 | 0.0117 | -0.0036 | -0.0411
| SP-18 | 0 | -0.01 | 0.01 | 42.6689 | -30.0612 | -20.0414 | -0.0381 | 0.0112 | -0.0034 | -0.0399
| SP-19 | 0 | -0.01 | 0.01 | 42.6078 | -30.3481 | -20.2989 | -0.0363 | 0.0105 | -0.0031 | -0.0379
| SP-20 | 0 | -0.01 | 0.01 | 42.5476 | -30.6178 | -20.5062 | -0.0336 | 0.0097 | -0.0028 | -0.0351
| SP-21 | 0 | -0.01 | 0.01 | 42.4827 | -30.2751 | -19.0338 | -0.066 | 0.0218 | -0.0099 | -0.0702
| SP-22 | 0 | -0.01 | 0.01 | 42.4126 | -30.597 | -19.2728 | -0.0682 | 0.0222 | -0.01 | -0.0724
| SP-23 | 0 | -0.01 | 0.01 | 42.3673 | -30.8514 | -19.5314 | -0.0513 | 0.0165 | -0.0072 | -0.0543
| SP-24 | 0 | -0.01 | 0.01 | 42.3114 | -31.1831 | -19.8958 | -0.0418 | 0.0131 | -0.0055 | -0.0441
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | -0.047017
| Std. deviation |  | 0.013354
| Max. Value |  | -0.029
| Min. Value |  | -0.0724
Inspection Group 74
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 42.1386 | -30.8407 | -18.2102 | -0.0489 | 0.0174 | -0.0104 | -0.053
| SP-14 | 0 | -0.01 | 0.01 | 42.0926 | -31.1028 | -18.4344 | -0.0645 | 0.0227 | -0.0133 | -0.0696
| SP-15 | 0 | -0.01 | 0.01 | 42.0044 | -31.4742 | -18.6391 | -0.059 | 0.0206 | -0.012 | -0.0636
| SP-16 | 0 | -0.01 | 0.01 | 41.9307 | -31.8251 | -18.8771 | -0.0471 | 0.0162 | -0.0094 | -0.0507
| SP-17 | 0 | -0.01 | 0.01 | 41.7841 | -31.2714 | -17.386 | -0.0086 | 0.0032 | -0.0023 | -0.0095
| SP-18 | 0 | -0.01 | 0.01 | 41.7268 | -31.6305 | -17.6768 | -0.0367 | 0.0135 | -0.0094 | -0.0402
| SP-19 | 0 | -0.01 | 0.01 | 41.6463 | -31.9505 | -17.8212 | -0.0436 | 0.016 | -0.0111 | -0.0478
| SP-20 | 0 | -0.01 | 0.01 | 41.6036 | -32.2522 | -18.0875 | -0.0136 | 0.0049 | -0.0034 | -0.0148
| SP-21 | 0 | -0.01 | 0.01 | 41.3626 | -31.6984 | -16.4935 | -0.0158 | 0.0061 | -0.0049 | -0.0176
| SP-22 | 0 | -0.01 | 0.01 | 41.2716 | -32.1036 | -16.6979 | -0.0383 | 0.0145 | -0.0118 | -0.0426
| SP-23 | 0 | -0.01 | 0.01 | 41.2129 | -32.4185 | -16.8943 | -0.0366 | 0.0137 | -0.0111 | -0.0407
| SP-24 | 0 | -0.01 | 0.01 | 41.1525 | -32.7619 | -17.1199 | -0.0152 | 0.0056 | -0.0045 | -0.0168
Number of points |  |  | 12
Out of Tolerance |  |  | 11
Performance |  |  | 8.333%
| Mean |  | -0.038908
| Std. deviation |  | 0.019945
| Max. Value |  | -0.0095
| Min. Value |  | -0.0696
Inspection Group 75
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 40.9151 | -32.1819 | -15.6967 | -0.0204 | 0.0078 | -0.007 | -0.0229
| SP-14 | 0 | -0.01 | 0.01 | 40.8599 | -32.4838 | -15.87 | -0.0238 | 0.0089 | -0.008 | -0.0267
| SP-15 | 0 | -0.01 | 0.01 | 40.8042 | -32.8487 | -16.1114 | -0.0125 | 0.0046 | -0.0041 | -0.014
| SP-16 | 0 | -0.01 | 0.01 | 40.7576 | -33.185 | -16.3486 | -0.0079 | 0.0029 | -0.0025 | -0.0088
| SP-17 | 0 | -0.01 | 0.01 | 40.473 | -32.5675 | -14.8404 | -0.0093 | 0.0034 | -0.0033 | -0.0104
| SP-18 | 0 | -0.01 | 0.01 | 40.42 | -32.8997 | -15.0329 | -0.0031 | 0.0011 | -0.0011 | -0.0034
| SP-19 | 0 | -0.01 | 0.01 | 40.3685 | -33.2431 | -15.2401 | 0.0264 | -0.0093 | 0.0089 | 0.0294
| SP-20 | 0 | -0.01 | 0.01 | 40.3119 | -33.6163 | -15.4619 | 0.0199 | -0.0069 | 0.0066 | 0.0221
| SP-21 | 0 | -0.01 | 0.01 | 40.0124 | -33.0403 | -13.9901 | -0.0211 | 0.0071 | -0.0072 | -0.0234
| SP-22 | 0 | -0.01 | 0.01 | 39.9647 | -33.4106 | -14.2137 | -0.0099 | 0.0032 | -0.0033 | -0.0109
| SP-23 | 0 | -0.01 | 0.01 | 39.9558 | -33.6752 | -14.4489 | 0.0074 | -0.0024 | 0.0024 | 0.0081
| SP-24 | 0 | -0.01 | 0.01 | 39.9205 | -34.0279 | -14.6913 | 0.0228 | -0.0073 | 0.0073 | 0.025
Number of points |  |  | 12
Out of Tolerance |  |  | 9
Performance |  |  | 25%
| Mean |  | -0.0029917
| Std. deviation |  | 0.019616
| Max. Value |  | 0.0294
| Min. Value |  | -0.0267
Inspection Group 76
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 39.598 | -33.5474 | -13.2093 | 0.017 | -0.005 | 0.0053 | 0.0185
| SP-14 | 0 | -0.01 | 0.01 | 39.5702 | -33.8778 | -13.4314 | 0.0119 | -0.0034 | 0.0036 | 0.0129
| SP-15 | 0 | -0.01 | 0.01 | 39.5329 | -34.1977 | -13.6098 | 0.0122 | -0.0034 | 0.0036 | 0.0132
| SP-16 | 0 | -0.01 | 0.01 | 39.5105 | -34.593 | -13.9053 | 0.0283 | -0.0077 | 0.0082 | 0.0305
| SP-17 | 0 | -0.01 | 0.01 | 39.2193 | -34.1362 | -12.4678 | 0.1011 | -0.0246 | 0.0281 | 0.1078
| SP-18 | 0 | -0.01 | 0.01 | 39.2272 | -34.4377 | -12.7595 | 0.0808 | -0.0192 | 0.022 | 0.0859
| SP-19 | 0 | -0.01 | 0.01 | 39.2353 | -34.6877 | -13.0067 | 0.0628 | -0.0146 | 0.0168 | 0.0666
| SP-20 | 0 | -0.01 | 0.01 | 39.2353 | -35.0283 | -13.3006 | 0.0749 | -0.0166 | 0.0194 | 0.0791
| SP-21 | 0 | -0.01 | 0.01 | 38.9304 | -34.7839 | -11.8413 | 0.1171 | -0.0183 | 0.0259 | 0.1213
| SP-22 | 0 | -0.01 | 0.01 | 38.9627 | -35.0261 | -12.1592 | 0.1063 | -0.0163 | 0.0232 | 0.11
| SP-23 | 0 | -0.01 | 0.01 | 38.9865 | -35.261 | -12.433 | 0.1019 | -0.015 | 0.0217 | 0.1052
| SP-24 | 0 | -0.01 | 0.01 | 39.0261 | -35.555 | -12.8221 | 0.1056 | -0.015 | 0.022 | 0.1089
Number of points |  |  | 12
Out of Tolerance |  |  | 12
Performance |  |  | 0%
| Mean |  | 0.071658
| Std. deviation |  | 0.042024
| Max. Value |  | 0.1213
| Min. Value |  | 0.0129
Inspection Group 77
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-1 | 0 | -0.01 | 0.01 | 38.7577 | -35.6106 | -11.2489 | 0.1859 | 0 | 0.0235 | 0.1873
| SP-2 | 0 | -0.01 | 0.01 | 38.7963 | -35.8666 | -11.5481 | 0.1609 | 0.0025 | 0.019 | 0.162
| SP-3 | 0 | -0.01 | 0.01 | 38.8312 | -36.0798 | -11.8123 | 0.1643 | 0.0048 | 0.0182 | 0.1654
| SP-4 | 0 | -0.01 | 0.01 | 38.8802 | -36.3491 | -12.1804 | 0.1446 | 0.0065 | 0.0147 | 0.1455
| SP-5 | 0 | -0.01 | 0.01 | 38.8172 | -36.4519 | -11.1764 | 0.1597 | 0.0257 | 0.0082 | 0.1619
| SP-6 | 0 | -0.01 | 0.01 | 38.8502 | -36.5763 | -11.4306 | 0.1686 | 0.028 | 0.0082 | 0.1711
| SP-7 | 0 | -0.01 | 0.01 | 38.9005 | -36.7373 | -11.9126 | 0.1616 | 0.026 | 0.0081 | 0.1639
| SP-8 | 0 | -0.01 | 0.01 | 38.887 | -36.839 | -10.6075 | 0.1458 | 0.0421 | 0.0013 | 0.1517
| SP-9 | 0 | -0.01 | 0.01 | 38.9194 | -36.943 | -10.8844 | 0.1312 | 0.0385 | 0.0009 | 0.1367
| SP-10 | 0 | -0.01 | 0.01 | 38.9601 | -37.0732 | -11.2315 | 0.1416 | 0.0426 | 0.0006 | 0.1479
| SP-11 | 0 | -0.01 | 0.01 | 39.0189 | -37.2609 | -11.6085 | 0.1438 | 0.0464 | -0.0003 | 0.1511
Number of points |  |  | 11
Out of Tolerance |  |  | 11
Performance |  |  | 0%
| Mean |  | 0.15859
| Std. deviation |  | 0.013951
| Max. Value |  | 0.1873
| Min. Value |  | 0.1367
Inspection Group 78
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-1 | 0 | -0.01 | 0.01 | 39.2776 | -37.7123 | -10.1675 | 0.0885 | 0.0537 | -0.006 | 0.1037
| SP-2 | 0 | -0.01 | 0.01 | 39.3301 | -37.8367 | -10.5173 | 0.0797 | 0.0506 | -0.0059 | 0.0947
| SP-3 | 0 | -0.01 | 0.01 | 39.3797 | -37.9486 | -10.8197 | 0.0748 | 0.0497 | -0.0059 | 0.09
| SP-4 | 0 | -0.01 | 0.01 | 39.4326 | -38.0632 | -11.1293 | 0.0729 | 0.051 | -0.0062 | 0.0892
| SP-5 | 0 | -0.01 | 0.01 | 39.6179 | -38.1646 | -9.9942 | 0.0515 | 0.0425 | -0.0052 | 0.0669
| SP-6 | 0 | -0.01 | 0.01 | 39.6316 | -38.2284 | -10.3786 | 0.0461 | 0.0385 | -0.0047 | 0.0602
| SP-7 | 0 | -0.01 | 0.01 | 39.6335 | -38.2754 | -10.7417 | 0.0408 | 0.0343 | -0.0042 | 0.0535
| SP-8 | 0 | -0.01 | 0.01 | 39.663 | -38.3618 | -11.1602 | 0.0626 | 0.0541 | -0.0067 | 0.083
| SP-9 | 0 | -0.01 | 0.01 | 39.9363 | -38.5127 | -9.9736 | 0.017 | 0.017 | -0.0019 | 0.0241
| SP-10 | 0 | -0.01 | 0.01 | 39.9518 | -38.5655 | -10.303 | 0.0045 | 0.0046 | -0.0005 | 0.0065
| SP-11 | 0 | -0.01 | 0.01 | 39.9719 | -38.6269 | -10.6748 | -0.0059 | -0.0061 | 0.0007 | -0.0085
| SP-12 | 0 | -0.01 | 0.01 | 39.9706 | -38.679 | -11.1491 | 0.0271 | 0.0289 | -0.0032 | 0.0398
Number of points |  |  | 12
Out of Tolerance |  |  | 10
Performance |  |  | 16.67%
| Mean |  | 0.058592
| Std. deviation |  | 0.036572
| Max. Value |  | 0.1037
| Min. Value |  | -0.0085
Inspection Group 79
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 40.8894 | -39.2836 | -9.8005 | -0.0176 | -0.0256 | 0.0017 | -0.0311
| SP-14 | 0 | -0.01 | 0.01 | 40.8916 | -39.3085 | -10.1469 | -0.0304 | -0.0454 | 0.003 | -0.0547
| SP-15 | 0 | -0.01 | 0.01 | 40.8626 | -39.3133 | -10.5072 | -0.0246 | -0.0374 | 0.0025 | -0.0448
| SP-16 | 0 | -0.01 | 0.01 | 40.9104 | -39.3711 | -10.9275 | -0.0112 | -0.0179 | 0.0011 | -0.0211
| SP-17 | 0 | -0.01 | 0.01 | 41.3966 | -39.6023 | -9.7542 | -0.0156 | -0.0266 | 0.0012 | -0.0309
| SP-18 | 0 | -0.01 | 0.01 | 41.3477 | -39.5915 | -10.1416 | -0.0273 | -0.0473 | 0.0022 | -0.0547
| SP-19 | 0 | -0.01 | 0.01 | 41.326 | -39.5941 | -10.4721 | -0.0172 | -0.0304 | 0.0014 | -0.0349
| SP-20 | 0 | -0.01 | 0.01 | 41.2881 | -39.5917 | -10.8942 | -0.0078 | -0.0142 | 0.0006 | -0.0162
| SP-21 | 0 | -0.01 | 0.01 | 42.0905 | -39.9692 | -9.7538 | -0.0207 | -0.0434 | 0.0007 | -0.0481
| SP-22 | 0 | -0.01 | 0.01 | 42.057 | -39.9593 | -10.1154 | -0.0204 | -0.0436 | 0.0007 | -0.0481
| SP-23 | 0 | -0.01 | 0.01 | 42.0582 | -39.9653 | -10.4675 | -0.014 | -0.0307 | 0.0004 | -0.0337
| SP-24 | 0 | -0.01 | 0.01 | 42.0441 | -39.9641 | -10.8416 | -0.0026 | -0.0059 | 0.0001 | -0.0065
Number of points |  |  | 12
Out of Tolerance |  |  | 11
Performance |  |  | 8.333%
| Mean |  | -0.0354
| Std. deviation |  | 0.015392
| Max. Value |  | -0.0065
| Min. Value |  | -0.0547
Inspection Group 80
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 43.0753 | -40.3837 | -9.6986 | -0.0116 | -0.0314 | -0.0005 | -0.0335
| SP-14 | 0 | -0.01 | 0.01 | 43.0841 | -40.3814 | -10.0265 | -0.0117 | -0.0326 | -0.0006 | -0.0347
| SP-15 | 0 | -0.01 | 0.01 | 43.0897 | -40.3757 | -10.4404 | -0.0101 | -0.0289 | -0.0006 | -0.0306
| SP-16 | 0 | -0.01 | 0.01 | 43.0567 | -40.3565 | -10.8405 | -0.0063 | -0.0184 | -0.0004 | -0.0194
| SP-17 | 0 | -0.01 | 0.01 | 43.8741 | -40.6523 | -9.699 | -0.0045 | -0.0148 | -0.0005 | -0.0154
| SP-18 | 0 | -0.01 | 0.01 | 43.8704 | -40.6376 | -10.075 | -0.0043 | -0.0145 | -0.0005 | -0.0152
| SP-19 | 0 | -0.01 | 0.01 | 43.813 | -40.605 | -10.4987 | -0.0041 | -0.014 | -0.0005 | -0.0146
| SP-20 | 0 | -0.01 | 0.01 | 43.8149 | -40.5956 | -10.7657 | 0.0021 | 0.0072 | 0.0003 | 0.0075
| SP-21 | 0 | -0.01 | 0.01 | 44.7307 | -40.8953 | -9.6552 | -0.0135 | -0.052 | -0.0028 | -0.0538
| SP-22 | 0 | -0.01 | 0.01 | 44.7068 | -40.8633 | -10.1328 | -0.0081 | -0.0322 | -0.0018 | -0.0333
| SP-23 | 0 | -0.01 | 0.01 | 44.6833 | -40.8333 | -10.5721 | -0.0052 | -0.0215 | -0.0012 | -0.0222
| SP-24 | 0 | -0.01 | 0.01 | 44.7344 | -40.8277 | -10.8862 | 0.0009 | 0.0039 | 0.0002 | 0.004
Number of points |  |  | 12
Out of Tolerance |  |  | 10
Performance |  |  | 16.67%
| Mean |  | -0.021767
| Std. deviation |  | 0.017065
| Max. Value |  | 0.0075
| Min. Value |  | -0.0538
Inspection Group 81
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 45.6908 | -41.1151 | -9.7619 | -0.0098 | -0.0458 | -0.0034 | -0.0469
| SP-14 | 0 | -0.01 | 0.01 | 45.679 | -41.0888 | -10.0801 | -0.0086 | -0.0413 | -0.0031 | -0.0423
| SP-15 | 0 | -0.01 | 0.01 | 45.663 | -41.0579 | -10.4457 | -0.0047 | -0.0234 | -0.0018 | -0.0239
| SP-16 | 0 | -0.01 | 0.01 | 45.6418 | -41.025 | -10.8198 | -0.0008 | -0.0042 | -0.0003 | -0.0043
| SP-17 | 0 | -0.01 | 0.01 | 46.6983 | -41.3127 | -9.7512 | -0.0026 | -0.0149 | -0.0014 | -0.0152
| SP-18 | 0 | -0.01 | 0.01 | 46.7725 | -41.2852 | -10.1604 | -0.0041 | -0.0251 | -0.0025 | -0.0256
| SP-19 | 0 | -0.01 | 0.01 | 46.7575 | -41.2461 | -10.5278 | -0.0008 | -0.005 | -0.0005 | -0.0051
| SP-20 | 0 | -0.01 | 0.01 | 46.6934 | -41.2049 | -10.8402 | 0.0013 | 0.0084 | 0.0008 | 0.0086
| SP-21 | 0 | -0.01 | 0.01 | 47.5316 | -41.4369 | -9.8377 | -0.006 | -0.0418 | -0.0048 | -0.0425
| SP-22 | 0 | -0.01 | 0.01 | 47.4909 | -41.3998 | -10.1107 | -0.006 | -0.0429 | -0.0049 | -0.0436
| SP-23 | 0 | -0.01 | 0.01 | 47.4316 | -41.3454 | -10.5145 | -0.0028 | -0.0212 | -0.0024 | -0.0216
| SP-24 | 0 | -0.01 | 0.01 | 47.4044 | -41.3008 | -10.8721 | -0.0007 | -0.0058 | -0.0007 | -0.0058
Number of points |  |  | 12
Out of Tolerance |  |  | 8
Performance |  |  | 33.33%
| Mean |  | -0.02235
| Std. deviation |  | 0.018533
| Max. Value |  | 0.0086
| Min. Value |  | -0.0469
Inspection Group 82
Surface Points
| Name | Offset | -Tol. | +Tol. | X | Y | Z | dX | dY | dZ | DL
| SP-13 | 0 | -0.01 | 0.01 | 48.5273 | -41.5672 | -9.8105 | -0.0029 | -0.0257 | -0.0035 | -0.0261
| SP-14 | 0 | -0.01 | 0.01 | 48.4766 | -41.5111 | -10.1869 | -0.0032 | -0.0305 | -0.0041 | -0.0309
| SP-15 | 0 | -0.01 | 0.01 | 48.4311 | -41.4653 | -10.493 | -0.0026 | -0.0257 | -0.0035 | -0.0261
| SP-16 | 0 | -0.01 | 0.01 | 48.3821 | -41.4105 | -10.8651 | -0.001 | -0.0106 | -0.0014 | -0.0107
| SP-17 | 0 | -0.01 | 0.01 | 49.203 | -41.6425 | -9.7581 | 0.0008 | 0.0087 | 0.0013 | 0.0089
| SP-18 | 0 | -0.01 | 0.01 | 49.2357 | -41.5838 | -10.1763 | 0.0001 | 0.0013 | 0.0002 | 0.0013
| SP-19 | 0 | -0.01 | 0.01 | 49.2105 | -41.5313 | -10.518 | 0.0006 | 0.0074 | 0.0011 | 0.0075
| SP-20 | 0 | -0.01 | 0.01 | 49.1099 | -41.467 | -10.9021 | 0.001 | 0.0135 | 0.002 | 0.0136
| SP-21 | 0 | -0.01 | 0.01 | 50.257 | -41.7062 | -9.8497 | 0.0003 | 0.0049 | 0.0008 | 0.005
| SP-22 | 0 | -0.01 | 0.01 | 50.2517 | -41.6445 | -10.2243 | 0.0004 | 0.0079 | 0.0013 | 0.008
| SP-23 | 0 | -0.01 | 0.01 | 50.1883 | -41.585 | -10.5677 | 0.0003 | 0.0059 | 0.001 | 0.006
| SP-24 | 0 | -0.01 | 0.01 | 50.1073 | -41.5232 | -10.9219 | 0.0011 | 0.0243 | 0.004 | 0.0247
Number of points |  |  | 12
Out of Tolerance |  |  | 6
Performance |  |  | 50%
| Mean |  | -0.0015667
| Std. deviation |  | 0.017734
| Max. Value |  | 0.0247
| Min. Value |  | -0.0309
Reameasure Conical seats
Conical Seat 2
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 76.7512 |  | 76.749 |  | -0.0022
Target point |  | Y | -0.01 | 0.01 | -37.0189 |  | -37.018 |  | 0.0009
|  | Z | -0.01 | 0.01 | -25.0018 |  | -25.0088 |  | -0.007
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0074 |  | 0.0074
Conical Seat 3
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 96.1272 |  | 96.1217 |  | -0.0055
Target point |  | Y | -0.01 | 0.01 | -15.6549 |  | -15.6614 |  | -0.0065
|  | Z | -0.01 | 0.01 | -24.4919 |  | -24.4924 |  | -0.0005
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0085 |  | 0.0085
Conical Seat 5
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 79.7088 |  | 79.7103 |  | 0.0015
Target point |  | Y | -0.01 | 0.01 | 30.9411 |  | 30.9435 |  | 0.0024
|  | Z | -0.01 | 0.01 | -27.5135 |  | -27.5059 |  | 0.0076
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0081 |  | 0.0081
Conical Seat 7
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 58.2322 |  | 58.2318 |  | -0.0004
Target point |  | Y | -0.01 | 0.01 | 48.5937 |  | 48.5976 |  | 0.0039
|  | Z | -0.01 | 0.01 | -9.7523 |  | -9.7506 |  | 0.0017
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0043 |  | 0.0043
Conical Seat 8
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 41.6065 |  | 41.6041 |  | -0.0024
Target point |  | Y | -0.01 | 0.01 | 49.9905 |  | 49.9946 |  | 0.0041
|  | Z | -0.01 | 0.01 | -6.2829 |  | -6.2835 |  | -0.0006
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0048 |  | 0.0048
Conical Seat 9
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 21.3958 |  | 21.3996 |  | 0.0038
Target point |  | Y | -0.01 | 0.01 | 29.1604 |  | 29.1592 |  | -0.0012
|  | Z | -0.01 | 0.01 | -3.7846 |  | -3.7893 |  | -0.0047
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0062 |  | 0.0062
Conical Seat 10
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 35.2895 |  | 35.2894 |  | -0.0001
Target point |  | Y | -0.01 | 0.01 | -2.2294 |  | -2.2318 |  | -0.0024
|  | Z | -0.01 | 0.01 | -8.9735 |  | -8.9793 |  | -0.0058
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0063 |  | 0.0063
Conical Seat 11
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 23.1781 |  | 23.1838 |  | 0.0057
Target point |  | Y | -0.01 | 0.01 | -25.0226 |  | -25.0252 |  | -0.0026
|  | Z | -0.01 | 0.01 | -2.602 |  | -2.5981 |  | 0.0039
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.0073 |  | 0.0073
Conical Seat 12
|  |  | -Tol. | +Tol. | Nominal |  | Measured |  | Dev/Mean |  | Error
|  | X | -0.01 | 0.01 | 34.0791 |  | 34.084 |  | 0.0049
Target point |  | Y | -0.01 | 0.01 | -45.8046 |  | -45.8043 |  | 0.0003
|  | Z | -0.01 | 0.01 | -3.0653 |  | -3.0619 |  | 0.0034
Distance To Target Point |  |  | -0.01 | 0.01 | 0 |  | 0.006 |  | 0.006