Major Tool & Machine, Inc.

Page: 1

1458 East 19th Street

MTM N/C: 19832

Date: 05/12/06

Indianapolis, IN 46218-4289

User ID: MCCORKLE

PRINCETON PLASMA PHYSICS LAB

Customer:

Telephone:

609-243-3655

Contact:

Mike Viola

mviola@pppl.gov

E-Mail:

609-243-2021

Fax:

S005243-F/Ln:2

Customer P.O.:

SE120-002 / PPPL NCSX VVSA

Part:

VVSA # 2

Serial No.:

1

Revision:

SE120-002

Drawing ID:

317-636-6433

Telephone:

DOUG MCCORKLE

Reported By:

Fax:

317-634-9420

E-Mail:

dMcCorkle@MajorTool.com

Problem:

During vacuum testing operation, the high vacuum gauge display would not illuminate (PPPL MKS Type 290

SP54-83 ION gauge controller).

A replacement gauge was not readily available.  After discussion with Mike Viola, it was decided to continue with

the test if we could verify the leak detector could accurately read the standard leak positioned at the opposite end of

the vessel.

All three low pressure gauges were pegged at zero for at least 14 hours prior to the test.  The turbo-molecular pump

had been running for approximately 8 hours prior to the test.  The MTM leak detector (Varian 979) was valved in,

and the PPPL roughing skid valved out and shut off.  Once the leak detector was ready, the test port pressure was

(0.0 x 10-4).  The MTM leak detector instantly recognized the standard leak located at the opposite end of the vessel

when activated.  After the standard leak helium was evacuated (waited approximately ½ hour), the leak test was

performed and revealed no leaks.  All welds and seals were sprayed, then the entire vessel was saturated.  No

response from the leak detector.

Proposed Disposition:

HOLD FOR CUSTOMER DISPOSITION

0

Number of additional pages:

Customer Disposition:       [   ] Use As Is          [  ] Rework          [  ] Repair          [  ] Scrap          [  ] Replace

Date:

Title:

Technical Contact Approval:

Date:

Title:

Buyer Approval:

Title:

Date:

Major Tool Implemented By:

n:\mtmapps\Mtnonc14.qrp

Major Tool and Machine, Inc.  1458 East 19th Street, Indianapolis, IN 46218-4289  Tel: 317-636-6433  Fax: 317-634-9420

Nonconformance Report: Major Tool NC19832

This is for:  VVSA # 2 Leak Check

Problem:

During vacuum testing operation, the high vacuum gauge display would not illuminate (PPPL MKS Type 290 SP54-83 ION gauge controller).

See further detail above regarding steps taken to verify validity of Vacuum leak check parameters.

The Major Tool Standard leak is 7.21E-8

Project Disposition:

Accept As Is

Approvals:

___________________________________________

Procurement Technical Representative

____________________________________________

Responsible Line Manager: