# Scope

# Refs

# Special Tools, Equipment and Materials

## Thermocouple Testing

### Digital Volt Meter with data logging capability

### Pass through connector (female side) with a 12” pigtail for each pin. Pigtail wire colors should be selected to simplify identifying thermocouple lead pairs for testing.

### Heat gun with operating heat range comparable to test range for Thermocouples.

### Small attachment to nose of heatgun that prevents variable distances from outlet of gun to the thermocouple during testing (see attached drawing).  This may have to be fabricated.

### Countdown timer with audible alarm to standardize exposure time of thermocouple to heat gun.

### Touchless thermometer or thermocouple attachment for DVM to record ambient temperature of Vessel thermocouples and calibrate heat gun.

## Heater Testing

### A test controller assembly that is capable of delivering power up to and including the maximum output required for operation.

#### It appears that Ray G. had build such a device for testing purposes.  Device has the ability to modify on/off ratio for variable heating level testing

# Pre-Reqs

# Test Procedures

## Thermocouples (TC)

### Tests should be performed in the following order.  Failure of any test procedure precludes further testing:

#### 1. Test for continuity between TC leads and Vessel assembly (ground)

##### Attach mating plug for pass-through (with pigtail)

##### Securely attach one DVM lead to body of Vessel at a well grounded point near area being tested

##### For each TC, test both leads individually for ground continuity and record any resistance values found.

##### Wiggle Test

###### Wiggle test each wire at any fusion welded junctions

###### Wiggle test each wire individually where it emerges from inconel braided sheath to enter the pass-through connector

###### Optionally wiggle test the entire wire bundle for the duration of its length if a fault is found in previous steps and cannot be localized.

##### *Any lead found with ground continuity should be considered a failure for that TC.

#### 2. Test for shorted lead pairs

##### Attach mating plug for pass through (with pigtail)

##### For each TC lead pair, attach DVM leads and record resistance value at ambient temperature (should be constant in Test Cell)

##### Wiggle Test

###### Wiggle test each wire at any fusion welded junctions

###### Wiggle test each wire individually where it emerges from inconel braided sheath to enter the pass-through connector

###### Optionally wiggle test the entire wire bundle for the duration of its length if a fault is found in previous steps and cannot be localized.

##### *Any lead pair found with unusual or out of range resistance values should be considered a failure for that TC.

#### 3. Test for leads shorted to pass through connector body

##### Remove mating “pigtail” plug from pass-though connector

##### Securely attach one DVM lead to body of pass through connector

##### For each thermocouple, test each pin lead individually for continuity with pass-through body

###### Wiggle test each wire where it enters the pass-through connector

###### Gently test each pin for tightness while testing for continuity with pass-through body

##### *Any lead found having continuity with the pass-through body should be considered a failure for that TC.

##### Any Pass through connector pin found to be loose enough to cause plug insertion problems or shorts to pass-through body should be considered a failure for that TC

#### 4. Test for Temperature response

##### Attach mating plug for pass through (with pigtail)

##### For each TC, attach DVM with resistance logging capability enabled to lead pair

##### Verify ambient temperature resistance value from earlier testing

##### Apply heat from heat gun for (TBD) seconds

###### Attach apparatus to nose of gun that will maintain consistent distance from TC

###### Allow gun to reach operating temperature before applying heat to TC (verify with thermometer)

###### Start countdown timer and immediately aim gun at TC

###### Apply heat until audible beep sounds.

###### Turn off gun and place in a secure location where it can cool and is not in danger of being inadvertently touched

###### Stop logging resistance values and record maximum.

## Heaters

### Tests should be performed in the following order.  Failure of any test procedure precludes further testing:

#### 1. Test for continuity between Heater leads and Vessel assembly (ground)

##### Attach mating plug for pass-through (with pigtail)

##### Securely attach one DVM lead to body of Vessel at a well grounded point near area being tested

##### For each heater, test both leads individually for ground continuity and record any resistance values found.

##### Wiggle Test

###### Wiggle test each wire at any fusion welded junctions

###### Wiggle test each wire individually where it enters the pass-through connector

##### *Any lead found with ground continuity should be considered a failure for that heater.

#### 2. Test for shorted lead pairs

##### Attach mating plug for pass through (with pigtail)

##### For each heater lead pair, attach DVM leads and record resistance value at ambient temperature (should be constant in Test Cell)

##### Wiggle Test

###### Wiggle test each wire at any fusion welded junctions

###### Wiggle test each wire individually where it emerges from inconel braided sheath to enter the pass-through connector

##### *Any lead pair found with unusual or out of range resistance values should be considered a failure for that heater.

#### 3. Test for leads shorted to pass through connector body

##### Remove mating “pigtail” plug from pass-though connector

##### Securely attach one DVM lead to body of pass through connector

##### For each thermocouple, test each pin lead individually for continuity with pass-through body

###### Wiggle test each wire where it enters the pass-through connector

###### Gently test each pin for tightness while testing for continuity with pass-through body

##### *Any lead found having continuity with the pass-through body should be considered a failure for that TC.

##### Any Pass through connector pin found to be loose enough to cause plug insertion problems or shorts to pass-through body should be considered a failure for that TC

#### 4. Test for output temperature

##### Remove mating pig-tail plug for pass through

##### For each heater, attach test device to heater pin pairs at pass-through connector (polarity is of no consequence)

##### Verify ambient temperature of heater assembly and record.

##### Start stopwatch and immediately plug test device into a 120v 60hz power outlet

##### Monitor heater temperature and stop stopwatch and unplug when heater reaches XX degrees.

##### Record time taken to reach operating temperature

## Flux Loops (per Ed, Stefan)

## Rogowski Coils (per Ed, Stefan)

## Voltage Loops (per Ed, Stefan)

# Conclusion

# Data Sheets