“Notes to My Successor”

Greg Tchilinguirian

July 22, 2008

Slight revision by Sichta 7-22-2008

(1) Scope & Deliverables (what was to be produced)

From Sichta/Tchilinguirian

Products/Deliverables:

- PTP for thermocouples and heaters.

  - Type-E MS-style thermocouple to 25 pin serial-D adaptor cable

  - MS- round connector to 25 pin serial-D connector adaptor cable

  - Heat gun attachment

  - Heater test box design and parts list

  - Testing data for thermocouples and heaters.

  - List of repaired leads

  - List of leads still needing repair

- PTP for Flux Loops.

  - Magnetic test device (J. E. Lawson)

  - Breakout test box (L. Guttadora)

- Rogowski Coil testing

- Voltage Loop testing

(2) Status (what was completed)

Some data sheets are in paper form, in Tchilinguirian’s WBS5 NCSX ‘paper’ folder. Files being maintained by WBS5 Manager, Computer Division.

Heater Testing -

VVSA1

- All testing finished finished except for heater temperature testing (no power applied).

VVSA2.

- Heater box J3 has been tested (fault was found).

- Heater box J4 has been wiggle tested.

VVSA3

- Wiggle testing has been completed for both J5 and J6

Thermocouple Testing –

VVSA1.

- Testing has been complete for each TC and connector.

VVSA2.

- Testing has been complete for each TC and connector.

VVSA3.

- All connectors and TC’s have been checked except for junction TC-J3-3.

- Pin to pin testing has been completed for TC-J3-3

Flux Loop Testing - PTP was completed pending review of involved parties.  Magnetic field test device has been completed.

Rogowski Coil testing - Rough version of PTP document exists.

Voltage Loop testing - None.

- Graphics or Photos – Included in .zip file:

DSC01301.JPG  Breakout Box. Full view

DSC01304.JPG  Breakout Box internal

DSC01308.JPG  Repaired/Spliced Thermocouple Leads

DSC01309.JPG  Type-E Thermocouple Connector – Port 12

DSC01311.JPG  Heater Tape attachment to Vacuum Vessel

DSC01324.JPG  Heater Connector -Port 12.  Thermocouple braid particles found

DSC01365.JPG  Heat-gun with standoff for thermocouple testing

DSC01368.JPG  Heater Tape performance testing.  Cart handle used as heat sink.

DSC01411.JPG  Damaged VVSA2 Type-E Thermocouple connector. (Repaired)

DSC01415.JPG  Flux Loop Test Device External

DSC01417.JPG  Flux Loop Test Device Internal

DSC01419.JPG  Damaged Heater lead junction Box Location

DSC01420.JPG  Damaged heater lead overview

DSC01421.JPG  Damaged heater lead at pass through

DSC01423.JPG  Damaged heater lead – damage location (at securing clip)

(4) Remaining scope -

Heater Testing -

VVSA1

- Heater temperature testing (no power applied) for J1 and J2.

VVSA2.

- Heater box J3 must be repaired and retested.

- Heater box J4 resistance testing has not been completed.

- Temperature testing has not been completed (no power applied) for J3 or J4.

VVSA3

- Resistance testing for J5 and J6.

- Temperature testing has not been completed (no power applied) for J5 or J6.

Thermocouple Testing –

- VVSA3 needs to have Junction TC-J3-3 tested.

Flux Loop Testing

- No testing has been conducted.

Rogowski Coil testing

- No testing has been conducted.

Voltage Loop testing

- No PTP document has been created, no testing has been conducted.

(5) Open issues or concerns

Heaters:

- Heater Leads run too close to heater tape.  There is a possibility that the leads will become damaged over time due to excessive heat exposure.

- Heater lead relieving is not tension free on VVSA2 J4.  The potential exists for insulation damage where the clips secure the leads to the vessel.

(6) List of documentation –

PTP’s for VV instrumentation testing were distributed for review but were never collected.  PTP numbers were assigned by ops center.

- Intellectual Property:

NONE

- Best Practices / Lessons Learned

NONE