| NATIONAL COMPACT STELLARATOR PROJECT Engineering Change Proposal (ECP) |  |  |  |  |
| --- | --- | --- | --- | --- |
| COVER PAGE (TO BE COMPLETED BY SYSTEMS ENGINEERING SUPPORT MANAGER) |  |  |  |  |
| Originator: Bob Simmons |  |  | Date: October 4, 2007 |  |
| ECP No: 058 |  | ECP Title: Update of FPA Station 1 CSPEC |  |  |
| Required Reviewers |  |  |  |  |
| Required Reviewers for this ECP: J.L. Anderson, H. Neilson, P. Heitzenroeder, P. Goranson, R. Ramakrishnan, P. Sichta, M. Viola, L. Dudek, C. Phelps, J. Malsbery, J. Levine, M. Zarnstorff |  |  |  |  |
| ECP Approval Level Expedited ECP? Yes No Change Level: Approving Official: |  |  |  |  |
| Actions - Update NCSX-CSPEC-185-01 (Station 1 Field Period Assembly) by 10/12/2007 - Update D-NCSX-FPA-001 (Field Period Assembly Station 1) by 10/12/2007 |  |  |  |  |
| APPROVALS (TO BE COMPLETED BY APPROVING OFFICIALS) |  |  |  |  |
| Change Level | Approving Official |  | Approval? | Signature |
| 3 | NCSX Project Manager |  | Yes No |  |
| 3a (Expedited ECP) | NCSX Engineering Manager |  | Yes No |  |
| 2 | NCSX Federal Project Director |  | Yes No |  |
| 1 | Associate Director OFES |  | Yes No |  |
| 0 | Deputy Secretary of Energy |  | Yes No |  |

| NATIONAL COMPACT STELLARATOR PROJECT Engineering Change Proposal (ECP) |  |
| --- | --- |
| PART I (TO BE COMPLETED BY ORIGINATOR) ECP- 058 |  |
| Originator: Bob Simmons | Date: October 4, 2007 |
| Overview of Change Type of ECP: EXPEDITED STANDARD Type of Change: TECHNICAL COST SCHEDULE EDITORIAL (Check all that Apply) Reason for Change: - Field measurements of thermocouples on VVSA segments did not meet either the CSPEC or Manufacturer’s requirements. Majority of longer thermocouples did meet requirements, but not the shorter ones. - Clarify pressure range and pressure flow requirements. Impacted WBS Elements: WBS 12 & WBS 18 Impacts of Change (Briefly Describe): - Thermocouples already installed on all three VVSA segments. Infeasible to remove/replace. - Original requirements would have necessitated purchase of expensive components and test equipment, when original requirement was only to check for blockages. These revised requirements will satisfy that need. Does this Change Impact Material Already Procured or Parts/Assemblies Already Assembled/Manufactured using this Material: Yes No If “Yes”, what is the recommended disposition of this material/part/assembly? “Use as is” – thermocouples already installed on all three VVSA segments. Assessment of Other Options: Originally, NCR 3719 was dispositioned (July 2007) to return to vendor. This is not feasible since all thermocouples installed on VVSAs. Revised NCR 3719 will show measurement ranges and disposition as “use as is” with a revised requirement. |  |

| NATIONAL COMPACT STELLARATOR PROJECT Engineering Change Proposal (ECP) |  |
| --- | --- |
| PART I (TO BE COMPLETED BY ORIGINATOR) ECP- 058 |  |
| Originator: Bob Simmons | Date: October 4, 2007 |
| Detailed Description of the Change: (Use Continuation Sheets and/or Attach Information/Sketches, As Needed) List Attachments, Impacted Documents, etc. - Updated NCSX-CSPEC-185-01 - Updated NCR 3719 - Updated D-NCSX- FPA-001 Description of Change: - Thermocouples: - Currently, both CSPEC 185-01-01, paragraphs 4.2.2.1.2 and 4.2.51, and D-NCSX-FPA-001-01, paragraph 6.12.2, specify the thermocouple electrical isolation requirements as >1.5 MOhms at 500V. This is in agreement with the Manufacturer’s specification. - Actual testing of the installed thermocouples on all three VVSA’s indicate that, in many instances, these requirements were not met. Specifically, the testing indicated: - VVSA #1: Measurements ranged between 0.60 – 1.40 MOhms at 600V (long thermocouples) - VVSA #1: Measurements ranged between 0.45 – 2.70 MOhms at 600V (short themocouples) - VVSA #2: Measurements ranged between 0.40 – 1.80 MOhms at 600V (long thermocouples) - VVSA #2: Measurements ranged between 1.60 – 2.80 MOhms at 600V (short themocouples) - VVSA #3: Measurements ranged between 0.43 – 0.80 MOhms at 500V (long thermocouples) - VVSA #3: Measurements ranged between 1.00 – 2.00 MOhms at 500V (short themocouples) - Since thermocouples are already installed, it is not feasible to remove and replace these thermocouples. - DO NOT CHANGE THE REQUIREMENT FROM THE VENDOR’S SPECIFICATION, BUT ADD A NOTE THAT STATES: “Revised NCR 3719 dispositioned nonconformances to “use as is” and installation procedure (D-NCSX-FPA-001 revised to permit readings of >0.4 MOhms at a minimum of 500V.” - Pressure range and Pressure-Flow Requirements: Change CSPEC 185-01-01, paragraph 4.2.1.1 to read: - “The heating/cooling system shall be leaked checked with helium gas at 300 psigin accordance with ENG-014.” - “The region around the compression fittings shall be checked with a helium leak detector tuned to 1 x 10-5 torr-l/s. Acceptance criteria shall be no detectable leaks.” |  |